# Mockup – D29447 6.1.11 Poor Film Seal (Inner vs Outer / POFS)

Purpose: Improve clarity by separating poor inner seal and poor outer seal (Outer film seal / POFS) while keeping 6.1.11 numbering unchanged.

Inspection technique (training note – no change to criteria): Use diffuse light (avoid direct spot light). View top-down, then tilt 10–20° to confirm lift/void vs glare. Glare moves with angle; a true lift/void tends to remain in place and may appear bright/milky (air gap).

## Table 11A – Poor inner seal (Inner film seal)

Scope: Sealed area around the port/chimney region (inner seal zone).

### Acceptance text (proposed):

Unacceptable: Any visible loss of adhesion (lift/void) in the defined seal area.

Acceptable: Seal is continuous and fully bonded around the circumference; no visible lift/void.

Optional (if acceptance criteria allow): “Acceptable – Observation (monitor)” for continuous and fully bonded but narrow/non-uniform appearance (no lift/void).

| Unacceptable / Minor Visual | Acceptable |
| --- | --- |
| INSERT PHOTO [Unacceptable Figure – insert image here] Caption (example): Visible lift/void in seal area. | INSERT PHOTO [Acceptable Figure – insert image here] Caption (example): Continuous, fully bonded seal. |
| Zone definition (text + optional diagram): Define the seal zone inspected for this table (inner vs outer). If helpful, add a simple top-view diagram with the zone highlighted. | Notes: Do not use seal width measurement; evaluate bonded vs not bonded. If in doubt, re-check with tilt (glare vs true lift). |
| — | Optional inset (inside Acceptable column): Acceptable – Observation (monitor) INSERT SMALL PHOTO (optional) Example caption: Narrow/non-uniform appearance but continuous and fully bonded; no lift/void. |

## Table 11B – Poor outer seal (Outer film seal / POFS)

Scope: Peripheral seal band near the lid edge (outer seal zone). “Poor outer seal” is referred to as POFS in PQC context.

### Acceptance text (proposed):

Unacceptable: Any visible loss of adhesion (lift/void) in the defined seal area.

Acceptable: Seal is continuous and fully bonded around the circumference; no visible lift/void.

Optional (if acceptance criteria allow): “Acceptable – Observation (monitor)” for continuous and fully bonded but narrow/non-uniform appearance (no lift/void).

| Unacceptable / Minor Visual | Acceptable |
| --- | --- |
| INSERT PHOTO [Unacceptable Figure – insert image here] Caption (example): Visible lift/void in seal area. | INSERT PHOTO [Acceptable Figure – insert image here] Caption (example): Continuous, fully bonded seal. |
| Zone definition (text + optional diagram): Define the seal zone inspected for this table (inner vs outer). If helpful, add a simple top-view diagram with the zone highlighted. | Notes: Do not use seal width measurement; evaluate bonded vs not bonded. If in doubt, re-check with tilt (glare vs true lift). |
| — | Optional inset (inside Acceptable column): Acceptable – Observation (monitor) INSERT SMALL PHOTO (optional) Example caption: Narrow/non-uniform appearance but continuous and fully bonded; no lift/void. |

## Cut-and-paste figure caption examples (optional):

Unacceptable Figure 11A-1 – Poor inner seal: Visible lift/void in inner seal zone.

Acceptable Figure 11A-2 – Good inner seal: Inner seal continuous and fully bonded.

Unacceptable Figure 11B-1 – Poor outer seal (POFS): Visible lift/void in outer seal zone.

Acceptable Figure 11B-2 – Good outer seal (POFS): Outer seal continuous and fully bonded.

Acceptable Figure 11B-3 – Outer seal observation (optional): Narrow/non-uniform but fully bonded; no lift/void.