## Dashboard
Click Less – Golden Standard Report (Template)
Run metadata
Assay
Assay Version
Reagent Lot(s)
Date range (Start Time)
# Instruments
# Modules
Input: Tests Summary
Input: Seal Test NEG
Input: Seal Test POS
Input: Worksheet
RuleBank Version
Engine Version
Triage counts | Bucket counts
FAIL | Major Functional Error
WARN | Minor Functional Error
INFO | Instrument Error
OK | Replacements
Sample ID or Test Type Deviation
Observation
Duplicates

## Information
Information
Skriptversion | v45.1.0
Användare | jesper.fredriksson
Datum Tid | 2026-01-07 08:16
Minitab Macro | %D12547-MTBU-SWE
CSV INFO
CSV | 2025 11 20 GXMTB RIF ULTRA 92303 Tests Summary.csv
LSP | 92303
Antal tester | 213
Dublett Sample ID | 1 (C15NOV23D_01_1_16+ x2)
Dublett Cartridge S/N | 2 (614334807 x2)
Använda INF/GX | Infinity-I; Infinity-V; Infinity-VI; Infinity-VIII
Bag Numbers Tested Using Infinity: | #00-10 x213
HEADER INFO
Worksheet | 2025 11 20 GXMTB RIF ULTRA 92303 Worksheet D25862.xlsx
Part No. | 700-5702 | Avvikande flik: Data - First 10
Batch No(s) | 1001506046 | Avvikande flik: Data - First 10
Cartridge No. (LSP) | 92303 | Avvikande flik: Data - First 10
Document Number | D25862 Attachment 1
Rev | AW
Effective | 9/11/2025
Seal Test POS | 2025 11 20 GXMTB RIF ULTRA 92303 Seal Test Pos.xlsx
Document Number | D10552
Rev | AU
Effective | 07/02/2025
Seal Test NEG | 2025 11 20 GXMTB RIF ULTRA 92303 Seal Test Neg.xlsx
Document Number | D10552
Rev | AU
Effective | 07/02/2025
LÄNKAR
SharePoint Batch Link | Öppna 1001506046
IPT App | LÄNK
MES Login | LÄNK
CSV Uploader | LÄNK
BMRAM | LÄNK
Agile | LÄNK
Note on A1: Generell spårbarhetsinfo samt Minitab-makro om tillgängligt.
Note on A7: Kort summering av vald .csv-fil.
Note on B14: Visar alltid Bag Number och summerar antal tester på Infinity.
Note on A16: Extreherar information från respektive fil och reagerar på avvikande innehåll.
Note on A33: Direktlänkar till funnet "Batch Number" i SharePoint (om valt) samt diverse IPT-relaterade verktyg.

## Run Findings
Run-level findings (file-level)
Severity | Bucket | RuleId | Title | Count | Evidence | SuggestedAction
FAIL | Instrument Error | INSTRUMENT_ERROR | Instrument Error (re-test/replacement needed) | 4
WARN | Minor Functional Error | MINOR_FUNCTIONAL_ERROR | Minor Functional Error (no re-test) | 4
WARN | Duplicates | DUPLICATE_SAMPLE_ID | Duplicate Sample ID detected | 2 | Lines 34, 35 for C15NOV23D_01_1_16+
WARN | Duplicates | DUPLICATE_CARTRIDGE_SN | Duplicate Cartridge S/N detected | 2 | Lines 42, 43 for 614334807
INFO | Replacements | REPLACEMENTS_PRESENT | Replacements detected (A/AA/AAA) | 8
INFO | Observation | UNKNOWN_ERROR_CODE | Observation: Error code not in RuleBank ErrorBank list. | 7 | Unknown codes: 5007, 2005, 1002, 1001, 5006, 2008
INFO | Observation | MISSING_BAG_SAMPLES | Missing samples by bag | 0 | Bag 1 is missing sample 17 Bag 5 is missing sample 11 Bag 5 is missing sample 12 Bag 6 is missing sample 20

## Information2
Triage view (row-level). Klicka på Row# för att hoppa till Raw.
Severity | Bucket | RuleId | Reason | SuggestedAction | Sample ID | Cartridge S/N | Test Type | Status | Test Result | ErrorCode | Error | Max Pressure (PSI) | Instrument S/N | Module S/N | Start Time | Row#
WARN | Minor Functional Error |  | MINOR_FUNCTIONAL_ERROR |  | B01NOV24B_03_2_19X | 614337470 | Positive Control 2 | Aborted | NO RESULT | 2037 | Error 2037: The cartridge integrity test failed at valve position 1390. The pressure change of -0.6 PSI did not exceed the requirement of 4.0 PSI. The pressure increased from 3.5 PSI to 3.0 PSI during the test | 20 | 802069 | 635193 | 2025-10-26T20:33:47+0100 | Go
WARN | Duplicates |  | DUPLICATE_SAMPLE_ID |  | C15NOV23D_01_1_16+ | 614332205 | Positive Control 1 | Done | MTB DETECTED LOW;RIF Resistance DETECTED |  |  | 15.5 | 839032 | 676886 | 2025-10-26T01:54:19+0200 | Go
WARN | Duplicates |  | DUPLICATE_SAMPLE_ID |  | C15NOV23D_01_1_16+ | 614332207 | Positive Control 1 | Done | MTB DETECTED LOW;RIF Resistance DETECTED |  |  | 20.3 | 839032 | 210024182 | 2025-10-26T01:54:46+0200 | Go
WARN | Minor Functional Error |  | MINOR_FUNCTIONAL_ERROR |  | C15NOV23D_02_1_12X | 614334122 | Positive Control 1 | Aborted | ERROR | 5007 | Error 5007: [CT1] probe check failed. Probe check value of 17 for reading number 2 was below the minimum of 110¦Error 5007: [SPC] probe check failed. Probe check value of 19 for reading number 2 was below the minimum of 119 | 17.4 | 839032 | 210058622 | 26/10/25 03:31:57 | Go
WARN | Duplicates |  | DUPLICATE_CARTRIDGE_SN |  | C15NOV23D_02_1_16+ | 614334807 | Positive Control 1 | Done | MTB DETECTED LOW;RIF Resistance DETECTED |  |  | 24.8 | 839032 | 782120 | 26/10/25 03:33:38 | Go
WARN | Duplicates |  | DUPLICATE_CARTRIDGE_SN |  | C15NOV23D_02_1_17+ | 614334807 | Positive Control 1 | Done | MTB DETECTED LOW;RIF Resistance DETECTED |  |  | 17.1 | 839032 | 676886 | 26/10/25 03:34:24 | Go
INFO | Replacements |  | REPLACEMENT_A |  | C15NOV23D_03_1_14AX | 614338044 | Positive Control 1 | Done | MTB DETECTED LOW;RIF Resistance DETECTED |  |  | 20 | 839032 | 656159 | 2025-10-26T08:58:18+0100 | Go
FAIL | Instrument Error |  | INSTRUMENT_ERROR |  | C15NOV23D_03_1_14X | 614338042 | Positive Control 1 | Aborted | ERROR | 2005 | Error 2005: Motion of the syringe drive was not detected. Detected motion started at position 250 uL and transferred 211 uL at valve position 245 with pressure 2.7 PSI | 19.2 | 839032 | 656159 | 2025-10-26T04:57:00+0100 | Go
INFO | Replacements |  | REPLACEMENT_A |  | C15NOV23D_03_1_15AX | 614338046 | Positive Control 1 | Done | MTB DETECTED LOW;RIF Resistance DETECTED |  |  | 24.6 | 839032 | 210058622 | 2025-10-26T08:59:18+0100 | Go
FAIL | Instrument Error |  | INSTRUMENT_ERROR |  | C15NOV23D_03_1_15X | 614338040 | Positive Control 1 | Done | ERROR | 1002 | Error 1002: The temperature difference of 5.3 ºC exceeds the limit of 5.0 ºC. The temperatures for heaters A and B are 97.3 ºC and 92.0 ºC | 24.6 | 839032 | 210059038 | 2025-10-26T04:57:26+0100 | Go
FAIL | Instrument Error |  | INSTRUMENT_ERROR |  | C15NOV23D_03_1_16+ | 614338465 | Positive Control 1 | Done | ERROR | 1001 | Error 1001: The actual temperature 61.6 ºC has drifted too far away from the setpoint of 61.0 ºC¦Error 1001: The actual temperature 60.5 ºC has drifted too far away from the setpoint of 61.0 ºC¦Error 1001: The actual temperature 60.4 ºC has drifted too far away from the setpoint of 61.0 ºC¦Error 1001: The actual temperature 61.7 ºC has drifted too far away from the setpoint of 61.0 ºC | 17.3 | 839032 | 210012379 | 2025-10-26T04:57:53+0100 | Go
INFO | Replacements |  | REPLACEMENT_A |  | C15NOV23D_03_1_16A+ | 614338467 | Positive Control 1 | Done | MTB DETECTED LOW;RIF Resistance DETECTED |  |  | 15.3 | 839032 | 652304 | 2025-10-26T09:00:18+0100 | Go
INFO | Replacements |  | REPLACEMENT_A |  | C15NOV23D_04_1_14AX | 614341834 | Positive Control 1 | Done | MTB DETECTED LOW;RIF Resistance DETECTED |  |  | 20.1 | 839032 | 210018949 | 2025-10-26T13:13:00+0100 | Go
INFO | Replacements |  | REPLACEMENT_A |  | C15NOV23D_06_1_13AAX | 614345588 | Positive Control 1 | Done | MTB DETECTED LOW;RIF Resistance DETECTED |  |  | 17.9 | 839032 | 690802 | 2025-10-26T17:29:12+0100 | Go
FAIL | Instrument Error |  | REPLACEMENT_A; INSTRUMENT_ERROR |  | C15NOV23D_06_1_13AX | 614345586 | Positive Control 1 | Aborted | ERROR | 2005 | Error 2005: Motion of the syringe drive was not detected. Detected motion started at position 250 uL and transferred 211 uL at valve position 245 with pressure 2.7 PSI | 18 | 839032 | 210018949 | 2025-10-26T13:14:16+0100 | Go
INFO | Replacements |  | REPLACEMENT_A |  | C15NOV23D_06_1_16A+_D06 | 614348455 | Positive Control 1 | Done | MTB DETECTED LOW;RIF Resistance DETECTED |  |  | 28.4 | 839032 | 210035802 | 2025-10-26T13:15:30+0100 | Go
WARN | Minor Functional Error |  | MINOR_FUNCTIONAL_ERROR |  | C15NOV23D_07_1_13X_D2B | 614349478 | Positive Control 1 | Aborted | ERROR | 5006 | Error 5006: [QC-2] probe check failed. Probe check value of 247 for reading number 1 was above the maximum of 223¦Error 5006: [QC-2] probe check failed. Probe check value of 247 for reading number 3 was above the maximum of 223 | 26 | 839032 | 730217 | 2025-10-26T15:01:05+0100 | Go
INFO | Replacements |  | REPLACEMENT_A |  | NEG_01_0_08A+ | 614332245 | Negative Control 1 | Done | MTB NOT DETECTED |  |  | 15.5 | 847922 | 730163 | 2025-10-26T02:34:57+0100 | Go
WARN | Minor Functional Error |  | MINOR_FUNCTIONAL_ERROR |  | NEG_10_0_06+ | 614390505 | Negative Control 1 | Aborted | ERROR | 2008 | Error 2008: Syringe pressure reading of 100.1 PSI exceeds the protocol limit of 100.0 PSI | 92 | 847922 | 701892 | 2025-10-26T23:42:50+0100 | Go

## Raw
Raw data (Tests Summary) + derived fields + tags
Assay | Assay Version | Sample ID | Cartridge S/N | Reagent Lot ID | Test Type | Instrument S/N | Module S/N | S/W Version | Start Time | Status | Test Result | Max Pressure (PSI) | Error | Prefix | Bag | ControlIndex | SampleNo2 | BagSampleRaw | ReplacementLevel | HasDelamination | SealID | SPC_Ct | ErrorCode | Tags | WorstBucket | Severity | _SourceExcelRow
Xpert MTB-RIF Ultra | 4 | B01NOV24B_01_2_19X | 614332206 | 92303 | Positive Control 2 | 802069 | 656295 | 7.1.6 | 2025-10-26T20:32:59+0100 | Done | MTB DETECTED LOW;RIF Resistance NOT DETECTED | 15.6 |  | B01NOV24B | 1 | 2 | 19 | 19X | 0 | FALSE |  |  |  |  | Observation | INFO | 10
Xpert MTB-RIF Ultra | 4 | B01NOV24B_01_2_20+ | 614332203 | 92303 | Positive Control 2 | 802069 | 624207 | 7.1.6 | 2025-10-26T20:36:56+0100 | Done | MTB DETECTED LOW;RIF Resistance NOT DETECTED | 20.8 |  | B01NOV24B | 1 | 2 | 20 | 20+ | 0 | FALSE |  |  |  |  | Observation | INFO | 11
Xpert MTB-RIF Ultra | 4 | B01NOV24B_02_2_19X | 614334366 | 92303 | Positive Control 2 | 802069 | 642168 | 7.1.6 | 2025-10-26T20:33:23+0100 | Done | MTB DETECTED LOW;RIF Resistance NOT DETECTED | 18.7 |  | B01NOV24B | 2 | 2 | 19 | 19X | 0 | FALSE |  |  |  |  | Observation | INFO | 12
Xpert MTB-RIF Ultra | 4 | B01NOV24B_02_2_20+ | 614334511 | 92303 | Positive Control 1 | 802069 | 775844 | 7.1.6 | 2025-10-26T20:37:22+0100 | Done | MTB DETECTED LOW;RIF Resistance NOT DETECTED | 21.9 |  | B01NOV24B | 2 | 2 | 20 | 20+ | 0 | FALSE |  |  |  |  | Observation | INFO | 13
Xpert MTB-RIF Ultra | 4 | B01NOV24B_03_2_19X | 614337470 | 92303 | Positive Control 2 | 802069 | 635193 | 7.1.6 | 2025-10-26T20:33:47+0100 | Aborted | NO RESULT | 20 | Error 2037: The cartridge integrity test failed at valve position 1390. The pressure change of -0.6 PSI did not exceed the requirement of 4.0 PSI. The pressure increased from 3.5 PSI to 3.0 PSI during the test | B01NOV24B | 3 | 2 | 19 | 19X | 0 | FALSE |  |  | 2037 | MINOR_FUNCTIONAL_ERROR | Minor Functional Error | WARN | 14
Xpert MTB-RIF Ultra | 4 | B01NOV24B_03_2_20+ | 614337317 | 92303 | Positive Control 2 | 802069 | 786669 | 7.1.6 | 2025-10-26T20:37:47+0100 | Done | MTB DETECTED LOW;RIF Resistance NOT DETECTED | 20.8 |  | B01NOV24B | 3 | 2 | 20 | 20+ | 0 | FALSE |  |  |  |  | Observation | INFO | 15
Xpert MTB-RIF Ultra | 4 | B01NOV24B_04_2_19X | 614341778 | 92303 | Positive Control 2 | 802069 | 659869 | 7.1.6 | 2025-10-26T20:34:11+0100 | Done | MTB DETECTED LOW;RIF Resistance NOT DETECTED | 19.7 |  | B01NOV24B | 4 | 2 | 19 | 19X | 0 | FALSE |  |  |  |  | Observation | INFO | 16
Xpert MTB-RIF Ultra | 4 | B01NOV24B_04_2_20+ | 614341609 | 92303 | Positive Control 2 | 802069 | 210052215 | 7.1.6 | 2025-10-26T20:38:11+0100 | Done | MTB DETECTED LOW;RIF Resistance NOT DETECTED | 20.9 |  | B01NOV24B | 4 | 2 | 20 | 20+ | 0 | FALSE |  |  |  |  | Observation | INFO | 17
Xpert MTB-RIF Ultra | 4 | B01NOV24B_05_2_19X | 614344026 | 92303 | Positive Control 2 | 802069 | 903819 | 7.1.6 | 2025-10-26T20:34:37+0100 | Done | MTB DETECTED LOW;RIF Resistance NOT DETECTED | 18.3 |  | B01NOV24B | 5 | 2 | 19 | 19X | 0 | FALSE |  |  |  |  | Observation | INFO | 18
Xpert MTB-RIF Ultra | 4 | B01NOV24B_05_2_20+ | 614345151 | 92303 | Positive Control 2 | 802069 | 805086 | 7.1.6 | 2025-10-26T20:38:38+0100 | Done | MTB DETECTED LOW;RIF Resistance NOT DETECTED | 21.3 |  | B01NOV24B | 5 | 2 | 20 | 20+ | 0 | FALSE |  |  |  |  | Observation | INFO | 19
Xpert MTB-RIF Ultra | 4 | B01NOV24B_06_2_19X | 614348164 | 92303 | Positive Control 2 | 802069 | 621821 | 7.1.6 | 2025-10-26T20:35:02+0100 | Done | MTB DETECTED LOW;RIF Resistance NOT DETECTED | 26.5 |  | B01NOV24B | 6 | 2 | 19 | 19X | 0 | FALSE |  |  |  |  | Observation | INFO | 20
Xpert MTB-RIF Ultra | 4 | B01NOV24B_07_2_19X | 614348958 | 92303 | Positive Control 2 | 802069 | 884165 | 7.1.6 | 2025-10-26T20:35:27+0100 | Done | MTB DETECTED LOW;RIF Resistance NOT DETECTED | 25.6 |  | B01NOV24B | 7 | 2 | 19 | 19X | 0 | FALSE |  |  |  |  | Observation | INFO | 21
Xpert MTB-RIF Ultra | 4 | B01NOV24B_07_2_20+ | 614349611 | 92303 | Positive Control 2 | 802069 | 869520 | 7.1.6 | 2025-10-26T20:39:04+0100 | Done | MTB DETECTED LOW;RIF Resistance NOT DETECTED | 30 |  | B01NOV24B | 7 | 2 | 20 | 20+ | 0 | FALSE |  |  |  |  | Observation | INFO | 22
Xpert MTB-RIF Ultra | 4 | B01NOV24B_08_2_19X | 614383062 | 92303 | Positive Control 2 | 802069 | 656453 | 7.1.6 | 2025-10-26T20:35:53+0100 | Done | MTB DETECTED LOW;RIF Resistance NOT DETECTED | 20 |  | B01NOV24B | 8 | 2 | 19 | 19X | 0 | FALSE |  |  |  |  | Observation | INFO | 23
Xpert MTB-RIF Ultra | 4 | B01NOV24B_08_2_20+ | 614382995 | 92303 | Positive Control 2 | 802069 | 719652 | 7.1.6 | 2025-10-26T20:39:28+0100 | Done | MTB DETECTED LOW;RIF Resistance NOT DETECTED | 23.7 |  | B01NOV24B | 8 | 2 | 20 | 20+ | 0 | FALSE |  |  |  |  | Observation | INFO | 24
Xpert MTB-RIF Ultra | 4 | B01NOV24B_09_2_19X | 614385874 | 92303 | Positive Control 2 | 802069 | 679144 | 7.1.6 | 2025-10-26T20:36:29+0100 | Done | MTB DETECTED LOW;RIF Resistance NOT DETECTED | 20.2 |  | B01NOV24B | 9 | 2 | 19 | 19X | 0 | FALSE |  |  |  |  | Observation | INFO | 25
Xpert MTB-RIF Ultra | 4 | B01NOV24B_09_2_20+ | 614386021 | 92303 | Positive Control 2 | 802069 | 686014 | 7.1.6 | 2025-10-26T20:39:52+0100 | Done | MTB DETECTED LOW;RIF Resistance NOT DETECTED | 20.5 |  | B01NOV24B | 9 | 2 | 20 | 20+ | 0 | FALSE |  |  |  |  | Observation | INFO | 26
Xpert MTB-RIF Ultra | 4 | B01NOV24B_10_2_19X | 614388340 | 92303 | Positive Control 2 | 839032 | 681294 | 6.8 | 2025-10-26T22:38:02+0100 | Done | MTB DETECTED LOW;RIF Resistance NOT DETECTED | 15.4 |  | B01NOV24B | 10 | 2 | 19 | 19X | 0 | FALSE |  |  |  |  | Observation | INFO | 27
Xpert MTB-RIF Ultra | 4 | B01NOV24B_10_2_20+ | 614388439 | 92303 | Positive Control 2 | 839032 | 210058622 | 6.8 | 2025-10-26T22:38:28+0100 | Done | MTB DETECTED LOW;RIF Resistance NOT DETECTED | 17.5 |  | B01NOV24B | 10 | 2 | 20 | 20+ | 0 | FALSE |  |  |  |  | Observation | INFO | 28
Xpert MTB-RIF Ultra | 4 | C15NOV23D_01_1_11X | 614332200 | 92303 | Positive Control 1 | 839032 | 652304 | 6.8 | 2025-10-26T01:52:14+0200 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 21.2 |  | C15NOV23D | 1 | 1 | 11 | 11X | 0 | FALSE |  |  |  |  | Observation | INFO | 29
Xpert MTB-RIF Ultra | 4 | C15NOV23D_01_1_12X | 614332198 | 92303 | Negative Control 1 | 839032 | 653259 | 6.8 | 2025-10-26T01:52:39+0200 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 19.8 |  | C15NOV23D | 1 | 1 | 12 | 12X | 0 | FALSE |  |  |  |  | Observation | INFO | 30
Xpert MTB-RIF Ultra | 4 | C15NOV23D_01_1_13X | 614332204 | 92303 | Positive Control 1 | 839032 | 627453 | 6.8 | 2025-10-26T01:53:06+0200 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 18.4 |  | C15NOV23D | 1 | 1 | 13 | 13X | 0 | FALSE |  |  |  |  | Observation | INFO | 31
Xpert MTB-RIF Ultra | 4 | C15NOV23D_01_1_14X | 614332202 | 92303 | Positive Control 1 | 839032 | 782120 | 6.8 | 2025-10-26T01:53:32+0200 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 24.8 |  | C15NOV23D | 1 | 1 | 14 | 14X | 0 | FALSE |  |  |  |  | Observation | INFO | 32
Xpert MTB-RIF Ultra | 4 | C15NOV23D_01_1_15+ | 614332201 | 92303 | Positive Control 1 | 839032 | 210034622 | 6.8 | 2025-10-26T01:53:57+0200 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 20.7 |  | C15NOV23D | 1 | 1 | 15 | 15+ | 0 | FALSE |  |  |  |  | Observation | INFO | 33
Xpert MTB-RIF Ultra | 4 | C15NOV23D_01_1_16+ | 614332205 | 92303 | Positive Control 1 | 839032 | 676886 | 6.8 | 2025-10-26T01:54:19+0200 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 15.5 |  | C15NOV23D | 1 | 1 | 16 | 16+ | 0 | FALSE |  |  |  | DUPLICATE_SAMPLE_ID | Duplicates | WARN | 34
Xpert MTB-RIF Ultra | 4 | C15NOV23D_01_1_16+ | 614332207 | 92303 | Positive Control 1 | 839032 | 210024182 | 6.8 | 2025-10-26T01:54:46+0200 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 20.3 |  | C15NOV23D | 1 | 1 | 16 | 16+ | 0 | FALSE |  |  |  | DUPLICATE_SAMPLE_ID | Duplicates | WARN | 35
Xpert MTB-RIF Ultra | 4 | C15NOV23D_01_1_18+ | 614332209 | 92303 | Positive Control 1 | 839032 | 210077690 | 6.8 | 2025-10-26T01:55:09+0200 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 22 |  | C15NOV23D | 1 | 1 | 18 | 18+ | 0 | FALSE |  |  |  |  | Observation | INFO | 36
Xpert MTB-RIF Ultra | 4 | C15NOV23D_02_1_11X | 614334666 | 92303 | Positive Control 1 | 839032 | 681294 | 6.8 | 26/10/25 03:31:32 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 18.3 |  | C15NOV23D | 2 | 1 | 11 | 11X | 0 | FALSE |  |  |  |  | Observation | INFO | 37
Xpert MTB-RIF Ultra | 4 | C15NOV23D_02_1_12X | 614334122 | 92303 | Positive Control 1 | 839032 | 210058622 | 6.8 | 26/10/25 03:31:57 | Aborted | ERROR | 17.4 | Error 5007: [CT1] probe check failed. Probe check value of 17 for reading number 2 was below the minimum of 110¦Error 5007: [SPC] probe check failed. Probe check value of 19 for reading number 2 was below the minimum of 119 | C15NOV23D | 2 | 1 | 12 | 12X | 0 | FALSE |  |  | 5007 | MINOR_FUNCTIONAL_ERROR | Minor Functional Error | WARN | 38
Xpert MTB-RIF Ultra | 4 | C15NOV23D_02_1_13X | 614334946 | 92303 | Positive Control 1 | 839032 | 652304 | 6.8 | 26/10/25 03:32:23 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 19.7 |  | C15NOV23D | 2 | 1 | 13 | 13X | 0 | FALSE |  |  |  |  | Observation | INFO | 39
Xpert MTB-RIF Ultra | 4 | C15NOV23D_02_1_14X | 614335244 | 92303 | Positive Control 1 | 839032 | 653259 | 6.8 | 26/10/25 03:32:47 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 19.2 |  | C15NOV23D | 2 | 1 | 14 | 14X | 0 | FALSE |  |  |  |  | Observation | INFO | 40
Xpert MTB-RIF Ultra | 4 | C15NOV23D_02_1_15+ | 614333907 | 92303 | Positive Control 1 | 839032 | 627453 | 6.8 | 26/10/25 03:33:13 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 15.8 |  | C15NOV23D | 2 | 1 | 15 | 15+ | 0 | FALSE |  |  |  |  | Observation | INFO | 41
Xpert MTB-RIF Ultra | 4 | C15NOV23D_02_1_16+ | 614334807 | 92303 | Positive Control 1 | 839032 | 782120 | 6.8 | 26/10/25 03:33:38 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 24.8 |  | C15NOV23D | 2 | 1 | 16 | 16+ | 0 | FALSE |  |  |  | DUPLICATE_CARTRIDGE_SN | Duplicates | WARN | 42
Xpert MTB-RIF Ultra | 4 | C15NOV23D_02_1_17+ | 614334807 | 92303 | Positive Control 1 | 839032 | 676886 | 6.8 | 26/10/25 03:34:24 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 17.1 |  | C15NOV23D | 2 | 1 | 17 | 17+ | 0 | FALSE |  |  |  | DUPLICATE_CARTRIDGE_SN | Duplicates | WARN | 43
Xpert MTB-RIF Ultra | 4 | C15NOV23D_02_1_18+ | 614334181 | 92303 | Positive Control 1 | 839032 | 210034622 | 6.8 | 26/10/25 03:34:03 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 19.6 |  | C15NOV23D | 2 | 1 | 18 | 18+ | 0 | FALSE |  |  |  |  | Observation | INFO | 44
Xpert MTB-RIF Ultra | 4 | C15NOV23D_03_1_11X | 614338636 | 92303 | Positive Control 1 | 839032 | 210016126 | 6.8 | 2025-10-26T04:53:46+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 16.7 |  | C15NOV23D | 3 | 1 | 11 | 11X | 0 | FALSE |  |  |  |  | Observation | INFO | 45
Xpert MTB-RIF Ultra | 4 | C15NOV23D_03_1_12X | 614337720 | 92303 | Positive Control 1 | 839032 | 736872 | 6.8 | 2025-10-26T04:55:01+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 21.1 |  | C15NOV23D | 3 | 1 | 12 | 12X | 0 | FALSE |  |  |  |  | Observation | INFO | 46
Xpert MTB-RIF Ultra | 4 | C15NOV23D_03_1_13X | 614338296 | 92303 | Positive Control 1 | 839032 | 739871 | 6.8 | 2025-10-26T04:56:13+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 19.3 |  | C15NOV23D | 3 | 1 | 13 | 13X | 0 | FALSE |  |  |  |  | Observation | INFO | 47
Xpert MTB-RIF Ultra | 4 | C15NOV23D_03_1_14AX | 614338044 | 92303 | Positive Control 1 | 839032 | 656159 | 6.8 | 2025-10-26T08:58:18+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 20 |  | C15NOV23D | 3 | 1 | 14 | 14AX | 1 | FALSE |  |  |  | REPLACEMENT_A | Replacements | INFO | 48
Xpert MTB-RIF Ultra | 4 | C15NOV23D_03_1_14X | 614338042 | 92303 | Positive Control 1 | 839032 | 656159 | 6.8 | 2025-10-26T04:57:00+0100 | Aborted | ERROR | 19.2 | Error 2005: Motion of the syringe drive was not detected. Detected motion started at position 250 uL and transferred 211 uL at valve position 245 with pressure 2.7 PSI | C15NOV23D | 3 | 1 | 14 | 14X | 0 | FALSE |  |  | 2005 | INSTRUMENT_ERROR | Instrument Error | FAIL | 49
Xpert MTB-RIF Ultra | 4 | C15NOV23D_03_1_15AX | 614338046 | 92303 | Positive Control 1 | 839032 | 210058622 | 6.8 | 2025-10-26T08:59:18+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 24.6 |  | C15NOV23D | 3 | 1 | 15 | 15AX | 1 | FALSE |  |  |  | REPLACEMENT_A | Replacements | INFO | 50
Xpert MTB-RIF Ultra | 4 | C15NOV23D_03_1_15X | 614338040 | 92303 | Positive Control 1 | 839032 | 210059038 | 6.8 | 2025-10-26T04:57:26+0100 | Done | ERROR | 24.6 | Error 1002: The temperature difference of 5.3 ºC exceeds the limit of 5.0 ºC. The temperatures for heaters A and B are 97.3 ºC and 92.0 ºC | C15NOV23D | 3 | 1 | 15 | 15X | 0 | FALSE |  |  | 1002 | INSTRUMENT_ERROR | Instrument Error | FAIL | 51
Xpert MTB-RIF Ultra | 4 | C15NOV23D_03_1_16+ | 614338465 | 92303 | Positive Control 1 | 839032 | 210012379 | 6.8 | 2025-10-26T04:57:53+0100 | Done | ERROR | 17.3 | Error 1001: The actual temperature 61.6 ºC has drifted too far away from the setpoint of 61.0 ºC¦Error 1001: The actual temperature 60.5 ºC has drifted too far away from the setpoint of 61.0 ºC¦Error 1001: The actual temperature 60.4 ºC has drifted too far away from the setpoint of 61.0 ºC¦Error 1001: The actual temperature 61.7 ºC has drifted too far away from the setpoint of 61.0 ºC | C15NOV23D | 3 | 1 | 16 | 16+ | 0 | FALSE |  |  | 1001 | INSTRUMENT_ERROR | Instrument Error | FAIL | 52
Xpert MTB-RIF Ultra | 4 | C15NOV23D_03_1_16A+ | 614338467 | 92303 | Positive Control 1 | 839032 | 652304 | 6.8 | 2025-10-26T09:00:18+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 15.3 |  | C15NOV23D | 3 | 1 | 16 | 16A+ | 1 | FALSE |  |  |  | REPLACEMENT_A | Replacements | INFO | 53
Xpert MTB-RIF Ultra | 4 | C15NOV23D_03_1_17+ | 614337607 | 92303 | Positive Control 1 | 839032 | 210003146 | 6.8 | 2025-10-26T04:58:16+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 19.5 |  | C15NOV23D | 3 | 1 | 17 | 17+ | 0 | FALSE |  |  |  |  | Observation | INFO | 54
Xpert MTB-RIF Ultra | 4 | C15NOV23D_03_1_18+ | 614338203 | 92303 | Positive Control 1 | 839032 | 694834 | 6.8 | 2025-10-26T04:58:39+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 16.8 |  | C15NOV23D | 3 | 1 | 18 | 18+ | 0 | FALSE |  |  |  |  | Observation | INFO | 55
Xpert MTB-RIF Ultra | 4 | C15NOV23D_04_1_11X | 614341522 | 92303 | Positive Control 1 | 839032 | 210016126 | 6.8 | 2025-10-26T08:52:27+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 15.3 |  | C15NOV23D | 4 | 1 | 11 | 11X | 0 | FALSE |  |  |  |  | Observation | INFO | 56
Xpert MTB-RIF Ultra | 4 | C15NOV23D_04_1_12X | 614341214 | 92303 | Positive Control 1 | 839032 | 736872 | 6.8 | 2025-10-26T08:52:53+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 21.1 |  | C15NOV23D | 4 | 1 | 12 | 12X | 0 | FALSE |  |  |  |  | Observation | INFO | 57
Xpert MTB-RIF Ultra | 4 | C15NOV23D_04_1_13X | 614340896 | 92303 | Positive Control 1 | 839032 | 739871 | 6.8 | 2025-10-26T08:53:17+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 18.4 |  | C15NOV23D | 4 | 1 | 13 | 13X | 0 | FALSE |  |  |  |  | Observation | INFO | 58
Xpert MTB-RIF Ultra | 4 | C15NOV23D_04_1_14AX | 614341834 | 92303 | Positive Control 1 | 839032 | 210018949 | 6.8 | 2025-10-26T13:13:00+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 20.1 |  | C15NOV23D | 4 | 1 | 14 | 14AX | 1 | FALSE |  |  |  | REPLACEMENT_A | Replacements | INFO | 59
Xpert MTB-RIF Ultra | 4 | C15NOV23D_04_1_15+ | 614341037 | 92303 | Positive Control 1 | 839032 | 210059038 | 6.8 | 2025-10-26T08:53:42+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 19.8 |  | C15NOV23D | 4 | 1 | 15 | 15+ | 0 | FALSE |  |  |  |  | Observation | INFO | 60
Xpert MTB-RIF Ultra | 4 | C15NOV23D_04_1_16+ | 614341409 | 92303 | Positive Control 1 | 839032 | 210012379 | 6.8 | 2025-10-26T08:54:10+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 18.2 |  | C15NOV23D | 4 | 1 | 16 | 16+ | 0 | FALSE |  |  |  |  | Observation | INFO | 61
Xpert MTB-RIF Ultra | 4 | C15NOV23D_04_1_17+ | 614340747 | 92303 | Positive Control 1 | 839032 | 210003146 | 6.8 | 2025-10-26T08:54:32+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 21.1 |  | C15NOV23D | 4 | 1 | 17 | 17+ | 0 | FALSE |  |  |  |  | Observation | INFO | 62
Xpert MTB-RIF Ultra | 4 | C15NOV23D_04_1_18+ | 614341905 | 92303 | Positive Control 1 | 839032 | 694834 | 6.8 | 2025-10-26T08:54:58+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 18.1 |  | C15NOV23D | 4 | 1 | 18 | 18+ | 0 | FALSE |  |  |  |  | Observation | INFO | 63
Xpert MTB-RIF Ultra | 4 | C15NOV23D_05_1_10X | 614344758 | 92303 | Positive Control 1 | 839032 | 681294 | 6.8 | 2025-10-26T08:55:24+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 18 |  | C15NOV23D | 5 | 1 | 10 | 10X | 0 | FALSE |  |  |  |  | Observation | INFO | 64
Xpert MTB-RIF Ultra | 4 | C15NOV23D_05_1_122+ | 614345044 | 92303 | Positive Control 1 | 839032 | 210058622 | 6.8 | 2025-10-26T08:55:49+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 18.7 |  | C15NOV23D | 5 | 1 | 122 | 122+ | 0 | FALSE |  |  |  |  | Observation | INFO | 65
Xpert MTB-RIF Ultra | 4 | C15NOV23D_05_1_13X | 614345292 | 92303 | Positive Control 1 | 839032 | 656159 | 6.8 | 2025-10-26T08:56:11+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 16.6 |  | C15NOV23D | 5 | 1 | 13 | 13X | 0 | FALSE |  |  |  |  | Observation | INFO | 66
Xpert MTB-RIF Ultra | 4 | C15NOV23D_05_1_14X | 614344398 | 92303 | Positive Control 1 | 839032 | 652304 | 6.8 | 2025-10-26T08:56:39+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 18 |  | C15NOV23D | 5 | 1 | 14 | 14X | 0 | FALSE |  |  |  |  | Observation | INFO | 67
Xpert MTB-RIF Ultra | 4 | C15NOV23D_05_1_15+ | 614344525 | 92303 | Positive Control 1 | 839032 | 653259 | 6.8 | 2025-10-26T08:57:03+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 17.6 |  | C15NOV23D | 5 | 1 | 15 | 15+ | 0 | FALSE |  |  |  |  | Observation | INFO | 68
Xpert MTB-RIF Ultra | 4 | C15NOV23D_05_1_16+ | 614344867 | 92303 | Positive Control 1 | 839032 | 627453 | 6.8 | 2025-10-26T08:57:29+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 18.1 |  | C15NOV23D | 5 | 1 | 16 | 16+ | 0 | FALSE |  |  |  |  | Observation | INFO | 69
Xpert MTB-RIF Ultra | 4 | C15NOV23D_05_1_17+ | 614344227 | 92303 | Positive Control 1 | 839032 | 676886 | 6.8 | 2025-10-26T08:57:53+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 16.5 |  | C15NOV23D | 5 | 1 | 17 | 17+ | 0 | FALSE |  |  |  |  | Observation | INFO | 70
Xpert MTB-RIF Ultra | 4 | C15NOV23D_05_1_18+ | 614343985 | 92303 | Positive Control 1 | 839032 | 210024182 | 6.8 | 2025-10-26T08:58:18+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 20 |  | C15NOV23D | 5 | 1 | 18 | 18+ | 0 | FALSE |  |  |  |  | Observation | INFO | 71
Xpert MTB-RIF Ultra | 4 | C15NOV23D_06_1_11X | 614348326 | 92303 | Positive Control 1 | 839032 | 210019897 | 6.8 | 2025-10-26T13:13:26+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 30.3 |  | C15NOV23D | 6 | 1 | 11 | 11X | 0 | FALSE |  |  |  |  | Observation | INFO | 72
Xpert MTB-RIF Ultra | 4 | C15NOV23D_06_1_12X | 614348666 | 92303 | Positive Control 1 | 839032 | 730217 | 6.8 | 2025-10-26T13:13:51+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 25.7 |  | C15NOV23D | 6 | 1 | 12 | 12X | 0 | FALSE |  |  |  |  | Observation | INFO | 73
Xpert MTB-RIF Ultra | 4 | C15NOV23D_06_1_13AAX | 614345588 | 92303 | Positive Control 1 | 839032 | 690802 | 6.8 | 2025-10-26T17:29:12+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 17.9 |  | C15NOV23D | 6 | 1 | 13 | 13AAX | 2 | FALSE |  |  |  | REPLACEMENT_A | Replacements | INFO | 74
Xpert MTB-RIF Ultra | 4 | C15NOV23D_06_1_13AX | 614345586 | 92303 | Positive Control 1 | 839032 | 210018949 | 6.8 | 2025-10-26T13:14:16+0100 | Aborted | ERROR | 18 | Error 2005: Motion of the syringe drive was not detected. Detected motion started at position 250 uL and transferred 211 uL at valve position 245 with pressure 2.7 PSI | C15NOV23D | 6 | 1 | 13 | 13AX | 1 | FALSE |  |  | 2005 | REPLACEMENT_A; INSTRUMENT_ERROR | Instrument Error | FAIL | 75
Xpert MTB-RIF Ultra | 4 | C15NOV23D_06_1_14X | 614346260 | 92303 | Positive Control 1 | 839032 | 210026044 | 6.8 | 2025-10-26T13:14:40+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 19 |  | C15NOV23D | 6 | 1 | 14 | 14X | 0 | FALSE |  |  |  |  | Observation | INFO | 76
Xpert MTB-RIF Ultra | 4 | C15NOV23D_06_1_15+ | 614345727 | 92303 | Positive Control 1 | 839032 | 638845 | 6.8 | 2025-10-26T13:15:05+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 19.4 |  | C15NOV23D | 6 | 1 | 15 | 15+ | 0 | FALSE |  |  |  |  | Observation | INFO | 77
Xpert MTB-RIF Ultra | 4 | C15NOV23D_06_1_16A+_D06 | 614348455 | 92303 | Positive Control 1 | 839032 | 210035802 | 6.8 | 2025-10-26T13:15:30+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 28.4 |  | C15NOV23D | 6 | 1 | 16 | 16A+ | 1 | FALSE |  |  |  | REPLACEMENT_A | Replacements | INFO | 78
Xpert MTB-RIF Ultra | 4 | C15NOV23D_06_1_17+ | 614347939 | 92303 | Positive Control 1 | 839032 | 880236 | 6.8 | 2025-10-26T13:15:56+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 26.7 |  | C15NOV23D | 6 | 1 | 17 | 17+ | 0 | FALSE |  |  |  |  | Observation | INFO | 79
Xpert MTB-RIF Ultra | 4 | C15NOV23D_06_1_18+ | 614346409 | 92303 | Positive Control 1 | 839032 | 210016126 | 6.8 | 2025-10-26T13:16:18+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 17.4 |  | C15NOV23D | 6 | 1 | 18 | 18+ | 0 | FALSE |  |  |  |  | Observation | INFO | 80
Xpert MTB-RIF Ultra | 4 | C15NOV23D_07_1_11X | 614350058 | 92303 | Positive Control 1 | 839032 | 210018949 | 6.8 | 2025-10-26T15:00:13+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 26.2 |  | C15NOV23D | 7 | 1 | 11 | 11X | 0 | FALSE |  |  |  |  | Observation | INFO | 81
Xpert MTB-RIF Ultra | 4 | C15NOV23D_07_1_12X | 614350324 | 92303 | Positive Control 1 | 839032 | 210019897 | 6.8 | 2025-10-26T15:00:39+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 24.4 |  | C15NOV23D | 7 | 1 | 12 | 12X | 0 | FALSE |  |  |  |  | Observation | INFO | 82
Xpert MTB-RIF Ultra | 4 | C15NOV23D_07_1_13X_D2B | 614349478 | 92303 | Positive Control 1 | 839032 | 730217 | 6.8 | 2025-10-26T15:01:05+0100 | Aborted | ERROR | 26 | Error 5006: [QC-2] probe check failed. Probe check value of 247 for reading number 1 was above the maximum of 223¦Error 5006: [QC-2] probe check failed. Probe check value of 247 for reading number 3 was above the maximum of 223 | C15NOV23D | 7 | 1 | 13 | 13X | 0 | FALSE |  |  | 5006 | MINOR_FUNCTIONAL_ERROR | Minor Functional Error | WARN | 83
Xpert MTB-RIF Ultra | 4 | C15NOV23D_07_1_14X | 614349732 | 92303 | Positive Control 1 | 839032 | 690802 | 6.8 | 2025-10-26T15:01:30+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 24 |  | C15NOV23D | 7 | 1 | 14 | 14X | 0 | FALSE |  |  |  |  | Observation | INFO | 84
Xpert MTB-RIF Ultra | 4 | C15NOV23D_07_1_15+ | 614349303 | 92303 | Positive Control 1 | 839032 | 210026044 | 6.8 | 2025-10-26T15:01:55+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 24.7 |  | C15NOV23D | 7 | 1 | 15 | 15+ | 0 | FALSE |  |  |  |  | Observation | INFO | 85
Xpert MTB-RIF Ultra | 4 | C15NOV23D_07_1_16+ | 614350191 | 92303 | Positive Control 1 | 839032 | 638845 | 6.8 | 2025-10-26T15:02:20+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 30.1 |  | C15NOV23D | 7 | 1 | 16 | 16+ | 0 | FALSE |  |  |  |  | Observation | INFO | 86
Xpert MTB-RIF Ultra | 4 | C15NOV23D_07_1_17+ | 614349857 | 92303 | Positive Control 1 | 839032 | 210035802 | 6.8 | 2025-10-26T15:02:46+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 22.7 |  | C15NOV23D | 7 | 1 | 17 | 17+ | 0 | FALSE |  |  |  |  | Observation | INFO | 87
Xpert MTB-RIF Ultra | 4 | C15NOV23D_07_1_18+ | 614349093 | 92303 | Positive Control 1 | 839032 | 880236 | 6.8 | 2025-10-26T15:03:10+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 31.3 |  | C15NOV23D | 7 | 1 | 18 | 18+ | 0 | FALSE |  |  |  |  | Observation | INFO | 88
Xpert MTB-RIF Ultra | 4 | C15NOV23D_08_1_11X | 614384828 | 92303 | Positive Control 1 | 839032 | 210018949 | 6.8 | 2025-10-26T17:29:12+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 17.6 |  | C15NOV23D | 8 | 1 | 11 | 11X | 0 | FALSE |  |  |  |  | Observation | INFO | 89
Xpert MTB-RIF Ultra | 4 | C15NOV23D_08_1_12X | 614384186 | 92303 | Positive Control 1 | 839032 | 210019897 | 6.8 | 2025-10-26T17:30:02+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 18.8 |  | C15NOV23D | 8 | 1 | 12 | 12X | 0 | FALSE |  |  |  |  | Observation | INFO | 90
Xpert MTB-RIF Ultra | 4 | C15NOV23D_08_1_13X | 614352594 | 92303 | Positive Control 1 | 839032 | 730217 | 6.8 | 2025-10-26T17:31:16+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 21.8 |  | C15NOV23D | 8 | 1 | 13 | 13X | 0 | FALSE |  |  |  |  | Observation | INFO | 91
Xpert MTB-RIF Ultra | 4 | C15NOV23D_08_1_14X_D2B | 614382746 | 92303 | Positive Control 1 | 839032 | 690802 | 6.8 | 2025-10-26T17:32:06+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 26.6 |  | C15NOV23D | 8 | 1 | 14 | 14X | 0 | FALSE |  |  |  |  | Observation | INFO | 92
Xpert MTB-RIF Ultra | 4 | C15NOV23D_08_1_15+ | 614383177 | 92303 | Positive Control 1 | 839032 | 210026044 | 6.8 | 2025-10-26T17:32:30+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 24.9 |  | C15NOV23D | 8 | 1 | 15 | 15+ | 0 | FALSE |  |  |  |  | Observation | INFO | 93
Xpert MTB-RIF Ultra | 4 | C15NOV23D_08_1_16+_D05 | 614384577 | 92303 | Positive Control 1 | 839032 | 638845 | 6.8 | 2025-10-26T17:32:55+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 19.7 |  | C15NOV23D | 8 | 1 | 16 | 16+ | 0 | FALSE |  |  |  |  | Observation | INFO | 94
Xpert MTB-RIF Ultra | 4 | C15NOV23D_08_1_17+ | 614352289 | 92303 | Positive Control 1 | 839032 | 210035802 | 6.8 | 2025-10-26T17:33:21+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 33.9 |  | C15NOV23D | 8 | 1 | 17 | 17+ | 0 | FALSE |  |  |  |  | Observation | INFO | 95
Xpert MTB-RIF Ultra | 4 | C15NOV23D_08_1_18+ | 614383633 | 92303 | Positive Control 1 | 839032 | 880236 | 6.8 | 2025-10-26T17:33:44+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 31.9 |  | C15NOV23D | 8 | 1 | 18 | 18+ | 0 | FALSE |  |  |  |  | Observation | INFO | 96
Xpert MTB-RIF Ultra | 4 | C15NOV23D_09_1_11X | 614385278 | 92303 | Positive Control 1 | 839032 | 210018949 | 6.8 | 2025-10-26T20:33:41+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 18.9 |  | C15NOV23D | 9 | 1 | 11 | 11X | 0 | FALSE |  |  |  |  | Observation | INFO | 97
Xpert MTB-RIF Ultra | 4 | C15NOV23D_09_1_12X | 614386394 | 92303 | Positive Control 1 | 839032 | 210019897 | 6.8 | 2025-10-26T20:34:06+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 19.6 |  | C15NOV23D | 9 | 1 | 12 | 12X | 0 | FALSE |  |  |  |  | Observation | INFO | 98
Xpert MTB-RIF Ultra | 4 | C15NOV23D_09_1_13X | 614385714 | 92303 | Positive Control 1 | 839032 | 730217 | 6.8 | 2025-10-26T20:34:32+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 19 |  | C15NOV23D | 9 | 1 | 13 | 13X | 0 | FALSE |  |  |  |  | Observation | INFO | 99
Xpert MTB-RIF Ultra | 4 | C15NOV23D_09_1_14X | 614386482 | 92303 | Positive Control 1 | 839032 | 690802 | 6.8 | 2025-10-26T20:34:58+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 17.8 |  | C15NOV23D | 9 | 1 | 14 | 14X | 0 | FALSE |  |  |  |  | Observation | INFO | 100
Xpert MTB-RIF Ultra | 4 | C15NOV23D_09_1_15+ | 614385059 | 92303 | Positive Control 1 | 839032 | 210026044 | 6.8 | 2025-10-26T20:35:22+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 17.5 |  | C15NOV23D | 9 | 1 | 15 | 15+ | 0 | FALSE |  |  |  |  | Observation | INFO | 101
Xpert MTB-RIF Ultra | 4 | C15NOV23D_09_1_16+ | 614385485 | 92303 | Positive Control 1 | 839032 | 638845 | 6.8 | 2025-10-26T20:35:47+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 19.5 |  | C15NOV23D | 9 | 1 | 16 | 16+ | 0 | FALSE |  |  |  |  | Observation | INFO | 102
Xpert MTB-RIF Ultra | 4 | C15NOV23D_09_1_17+ | 614386401 | 92303 | Positive Control 1 | 839032 | 210035802 | 6.8 | 2025-10-26T20:36:13+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 19.9 |  | C15NOV23D | 9 | 1 | 17 | 17+ | 0 | FALSE |  |  |  |  | Observation | INFO | 103
Xpert MTB-RIF Ultra | 4 | C15NOV23D_09_1_18+ | 614385629 | 92303 | Positive Control 1 | 839032 | 880236 | 6.8 | 2025-10-26T20:36:38+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 20.2 |  | C15NOV23D | 9 | 1 | 18 | 18+ | 0 | FALSE |  |  |  |  | Observation | INFO | 104
Xpert MTB-RIF Ultra | 4 | C15NOV23D_10_1_11X | 614388914 | 92303 | Positive Control 1 | 839032 | 210016126 | 6.8 | 2025-10-26T22:34:44+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 15.3 |  | C15NOV23D | 10 | 1 | 11 | 11X | 0 | FALSE |  |  |  |  | Observation | INFO | 105
Xpert MTB-RIF Ultra | 4 | C15NOV23D_10_1_12X | 614389210 | 92303 | Positive Control 1 | 839032 | 736872 | 6.8 | 2025-10-26T22:35:10+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 20.9 |  | C15NOV23D | 10 | 1 | 12 | 12X | 0 | FALSE |  |  |  |  | Observation | INFO | 106
Xpert MTB-RIF Ultra | 4 | C15NOV23D_10_1_13X | 614389594 | 92303 | Positive Control 1 | 839032 | 739871 | 6.8 | 2025-10-26T22:35:33+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 18.4 |  | C15NOV23D | 10 | 1 | 13 | 13X | 0 | FALSE |  |  |  |  | Observation | INFO | 107
Xpert MTB-RIF Ultra | 4 | C15NOV23D_10_1_14X | 614388652 | 92303 | Positive Control 1 | 839032 | 656159 | 6.8 | 2025-10-26T22:35:57+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 17.7 |  | C15NOV23D | 10 | 1 | 14 | 14X | 0 | FALSE |  |  |  |  | Observation | INFO | 108
Xpert MTB-RIF Ultra | 4 | C15NOV23D_10_1_15+ | 614389737 | 92303 | Positive Control 1 | 839032 | 210059038 | 6.8 | 2025-10-26T22:36:21+0100 | Done | MTB DETECTED LOW;RIF Resistance INDETERMINATE | 16.4 |  | C15NOV23D | 10 | 1 | 15 | 15+ | 0 | FALSE |  |  |  |  | Observation | INFO | 109
Xpert MTB-RIF Ultra | 4 | C15NOV23D_10_1_16+ | 614389057 | 92303 | Positive Control 1 | 839032 | 210012379 | 6.8 | 2025-10-26T22:36:46+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 18.3 |  | C15NOV23D | 10 | 1 | 16 | 16+ | 0 | FALSE |  |  |  |  | Observation | INFO | 110
Xpert MTB-RIF Ultra | 4 | C15NOV23D_10_1_17+ | 614389329 | 92303 | Positive Control 1 | 839032 | 210003146 | 6.8 | 2025-10-26T22:37:10+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 20.5 |  | C15NOV23D | 10 | 1 | 17 | 17+ | 0 | FALSE |  |  |  |  | Observation | INFO | 111
Xpert MTB-RIF Ultra | 4 | C15NOV23D_10_1_18+ | 614388817 | 92303 | Positive Control 1 | 839032 | 694834 | 6.8 | 2025-10-26T22:37:37+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 17.8 |  | C15NOV23D | 10 | 1 | 18 | 18+ | 0 | FALSE |  |  |  |  | Observation | INFO | 112
Xpert MTB-RIF Ultra | 4 | NEG_00_0_01X | 614332136 | 92303 | Negative Control 1 | 803094 | 730647 | 6.8 | 2025-10-26T01:27:08+0200 | Done | MTB NOT DETECTED | 18.9 |  | NEG | 0 | 0 | 01 | 01X | 0 | FALSE |  |  |  |  | Observation | INFO | 113
Xpert MTB-RIF Ultra | 4 | NEG_00_0_02X | 614332146 | 92303 | Negative Control 1 | 803094 | 773527 | 6.8 | 2025-10-26T01:27:35+0200 | Done | MTB NOT DETECTED | 19 |  | NEG | 0 | 0 | 02 | 02X | 0 | FALSE |  |  |  |  | Observation | INFO | 114
Xpert MTB-RIF Ultra | 4 | NEG_00_0_03X | 614332148 | 92303 | Negative Control 1 | 803094 | 210043140 | 6.8 | 2025-10-26T01:28:05+0200 | Done | MTB NOT DETECTED | 19.4 |  | NEG | 0 | 0 | 03 | 03X | 0 | FALSE |  |  |  |  | Observation | INFO | 115
Xpert MTB-RIF Ultra | 4 | NEG_00_0_04X | 614332134 | 92303 | Negative Control 1 | 803094 | 909441 | 6.8 | 2025-10-26T01:28:32+0200 | Done | MTB NOT DETECTED | 17.5 |  | NEG | 0 | 0 | 04 | 04X | 0 | FALSE |  |  |  |  | Observation | INFO | 116
Xpert MTB-RIF Ultra | 4 | NEG_00_0_05X | 614332132 | 92303 | Negative Control 1 | 803094 | 880243 | 6.8 | 2025-10-26T01:29:09+0200 | Done | MTB NOT DETECTED | 14.8 |  | NEG | 0 | 0 | 05 | 05X | 0 | FALSE |  |  |  |  | Observation | INFO | 117
Xpert MTB-RIF Ultra | 4 | NEG_00_0_06+ | 614332137 | 92303 | Negative Control 1 | 803094 | 210013102 | 6.8 | 2025-10-26T01:29:41+0200 | Done | MTB NOT DETECTED | 16.9 |  | NEG | 0 | 0 | 06 | 06+ | 0 | FALSE |  |  |  |  | Observation | INFO | 118
Xpert MTB-RIF Ultra | 4 | NEG_00_0_07+ | 614332149 | 92303 | Negative Control 1 | 803094 | 674169 | 6.8 | 2025-10-26T01:30:04+0200 | Done | MTB NOT DETECTED | 17.3 |  | NEG | 0 | 0 | 07 | 07+ | 0 | FALSE |  |  |  |  | Observation | INFO | 119
Xpert MTB-RIF Ultra | 4 | NEG_00_0_08+ | 614332145 | 92303 | Negative Control 1 | 803094 | 631549 | 6.8 | 2025-10-26T01:30:36+0200 | Done | MTB NOT DETECTED | 22.5 |  | NEG | 0 | 0 | 08 | 08+ | 0 | FALSE |  |  |  |  | Observation | INFO | 120
Xpert MTB-RIF Ultra | 4 | NEG_00_0_09+ | 614332151 | 92303 | Negative Control 1 | 803094 | 736909 | 6.8 | 2025-10-26T01:31:03+0200 | Done | MTB NOT DETECTED | 18.3 |  | NEG | 0 | 0 | 09 | 09+ | 0 | FALSE |  |  |  |  | Observation | INFO | 121
Xpert MTB-RIF Ultra | 4 | NEG_00_0_10+ | 614332135 | 92303 | Negative Control 1 | 803094 | 914633 | 6.8 | 2025-10-26T01:31:27+0200 | Done | MTB NOT DETECTED | 18.6 |  | NEG | 0 | 0 | 10 | 10+ | 0 | FALSE |  |  |  |  | Observation | INFO | 122
Xpert MTB-RIF Ultra | 4 | NEG_01_0_01X | 614332262 | 92303 | Negative Control 1 | 847922 | 604077 | 6.8 | 2025-10-26T02:56:07+0100 | Done | MTB NOT DETECTED | 19.2 |  | NEG | 1 | 0 | 01 | 01X | 0 | FALSE |  |  |  |  | Observation | INFO | 123
Xpert MTB-RIF Ultra | 4 | NEG_01_0_02X | 614332258 | 92303 | Negative Control 1 | 847922 | 210015588 | 6.8 | 2025-10-26T02:56:32+0100 | Done | MTB NOT DETECTED | 21.6 |  | NEG | 1 | 0 | 02 | 02X | 0 | FALSE |  |  |  |  | Observation | INFO | 124
Xpert MTB-RIF Ultra | 4 | NEG_01_0_03X | 614332256 | 92303 | Negative Control 1 | 847922 | 660764 | 6.8 | 2025-10-26T02:56:56+0100 | Done | MTB NOT DETECTED | 15.6 |  | NEG | 1 | 0 | 03 | 03X | 0 | FALSE |  |  |  |  | Observation | INFO | 125
Xpert MTB-RIF Ultra | 4 | NEG_01_0_04X | 614332260 | 92303 | Negative Control 1 | 847922 | 617583 | 6.8 | 2025-10-26T02:57:42+0100 | Done | MTB NOT DETECTED | 18.5 |  | NEG | 1 | 0 | 04 | 04X | 0 | FALSE |  |  |  |  | Observation | INFO | 126
Xpert MTB-RIF Ultra | 4 | NEG_01_0_05X | 614332264 | 92303 | Negative Control 1 | 847922 | 725382 | 6.8 | 2025-10-26T02:58:58+0100 | Done | MTB NOT DETECTED | 21.5 |  | NEG | 1 | 0 | 05 | 05X | 0 | FALSE |  |  |  |  | Observation | INFO | 127
Xpert MTB-RIF Ultra | 4 | NEG_01_0_06+ | 614332263 | 92303 | Negative Control 1 | 847922 | 710982 | 6.8 | 2025-10-26T02:59:48+0100 | Done | MTB NOT DETECTED | 22 |  | NEG | 1 | 0 | 06 | 06+ | 0 | FALSE |  |  |  |  | Observation | INFO | 128
Xpert MTB-RIF Ultra | 4 | NEG_01_0_07+ | 614332261 | 92303 | Negative Control 1 | 847922 | 912152 | 6.8 | 2025-10-26T02:00:11+0100 | Done | MTB NOT DETECTED | 16 |  | NEG | 1 | 0 | 07 | 07+ | 0 | FALSE |  |  |  |  | Observation | INFO | 129
Xpert MTB-RIF Ultra | 4 | NEG_01_0_08A+ | 614332245 | 92303 | Negative Control 1 | 847922 | 730163 | 6.8 | 2025-10-26T02:34:57+0100 | Done | MTB NOT DETECTED | 15.5 |  | NEG | 1 | 0 | 08 | 08A+ | 1 | FALSE |  |  |  | REPLACEMENT_A | Replacements | INFO | 130
Xpert MTB-RIF Ultra | 4 | NEG_01_0_09+ | 614332267 | 92303 | Negative Control 1 | 847922 | 892236 | 6.8 | 2025-10-26T02:00:36+0100 | Done | MTB NOT DETECTED | 18.9 |  | NEG | 1 | 0 | 09 | 09+ | 0 | FALSE |  |  |  |  | Observation | INFO | 131
Xpert MTB-RIF Ultra | 4 | NEG_01_0_10+ | 614332259 | 92303 | Negative Control 1 | 847922 | 210096147 | 6.8 | 2025-10-26T02:01:01+0100 | Done | MTB NOT DETECTED | 17.1 |  | NEG | 1 | 0 | 10 | 10+ | 0 | FALSE |  |  |  |  | Observation | INFO | 132
Xpert MTB-RIF Ultra | 4 | NEG_02_0_01X | 614333120 | 92303 | Negative Control 1 | 847922 | 897302 | 6.8 | 2025-10-26T03:39:14+0100 | Done | MTB NOT DETECTED | 17.2 |  | NEG | 2 | 0 | 01 | 01X | 0 | FALSE |  |  |  |  | Observation | INFO | 133
Xpert MTB-RIF Ultra | 4 | NEG_02_0_02X | 614332906 | 92303 | Negative Control 1 | 847922 | 210097747 | 6.8 | 2025-10-26T03:39:40+0100 | Done | MTB NOT DETECTED | 17 |  | NEG | 2 | 0 | 02 | 02X | 0 | FALSE |  |  |  |  | Observation | INFO | 134
Xpert MTB-RIF Ultra | 4 | NEG_02_0_03X | 614333486 | 92303 | Negative Control 1 | 847922 | 210013071 | 6.8 | 2025-10-26T03:40:05+0100 | Done | MTB NOT DETECTED | 19.6 |  | NEG | 2 | 0 | 03 | 03X | 0 | FALSE |  |  |  |  | Observation | INFO | 135
Xpert MTB-RIF Ultra | 4 | NEG_02_0_04X | 614333754 | 92303 | Negative Control 1 | 847922 | 657194 | 6.8 | 2025-10-26T03:40:54+0100 | Done | MTB NOT DETECTED | 21.1 |  | NEG | 2 | 0 | 04 | 04X | 0 | FALSE |  |  |  |  | Observation | INFO | 136
Xpert MTB-RIF Ultra | 4 | NEG_02_0_05X | 614332568 | 92303 | Negative Control 1 | 847922 | 702829 | 6.8 | 2025-10-26T03:41:19+0100 | Done | MTB NOT DETECTED | 17.6 |  | NEG | 2 | 0 | 05 | 05X | 0 | FALSE |  |  |  |  | Observation | INFO | 137
Xpert MTB-RIF Ultra | 4 | NEG_02_0_06+ | 614333339 | 92303 | Negative Control 1 | 847922 | 750444 | 6.8 | 2025-10-26T03:41:42+0100 | Done | MTB NOT DETECTED | 18.5 |  | NEG | 2 | 0 | 06 | 06+ | 0 | FALSE |  |  |  |  | Observation | INFO | 138
Xpert MTB-RIF Ultra | 4 | NEG_02_0_07+ | 614332989 | 92303 | Negative Control 1 | 847922 | 220000161 | 6.8 | 2025-10-26T03:42:06+0100 | Done | MTB NOT DETECTED | 19.8 |  | NEG | 2 | 0 | 07 | 07+ | 0 | FALSE |  |  |  |  | Observation | INFO | 139
Xpert MTB-RIF Ultra | 4 | NEG_02_0_08+ | 614332693 | 92303 | Negative Control 1 | 847922 | 760491 | 6.8 | 2025-10-26T03:42:31+0100 | Done | MTB NOT DETECTED | 19.9 |  | NEG | 2 | 0 | 08 | 08+ | 0 | FALSE |  |  |  |  | Observation | INFO | 140
Xpert MTB-RIF Ultra | 4 | NEG_02_0_09+ | 614333581 | 92303 | Negative Control 1 | 847922 | 210068154 | 6.8 | 2025-10-26T03:42:56+0100 | Done | MTB NOT DETECTED | 18.2 |  | NEG | 2 | 0 | 09 | 09+ | 0 | FALSE |  |  |  |  | Observation | INFO | 141
Xpert MTB-RIF Ultra | 4 | NEG_02_0_10+ | 614332391 | 92303 | Negative Control 1 | 847922 | 902140 | 6.8 | 2025-10-26T03:43:21+0100 | Done | MTB NOT DETECTED | 15.7 |  | NEG | 2 | 0 | 10 | 10+ | 0 | FALSE |  |  |  |  | Observation | INFO | 142
Xpert MTB-RIF Ultra | 4 | NEG_03_0_01X | 614337122 | 92303 | Negative Control 1 | 847922 | 897302 | 6.8 | 2025-10-26T05:52:14+0100 | Done | MTB NOT DETECTED | 12.7 |  | NEG | 3 | 0 | 01 | 01X | 0 | FALSE |  |  |  |  | Observation | INFO | 143
Xpert MTB-RIF Ultra | 4 | NEG_03_0_02X | 614335906 | 92303 | Negative Control 1 | 847922 | 210097747 | 6.8 | 2025-10-26T05:52:40+0100 | Done | MTB NOT DETECTED | 22.3 |  | NEG | 3 | 0 | 02 | 02X | 0 | FALSE |  |  |  |  | Observation | INFO | 144
Xpert MTB-RIF Ultra | 4 | NEG_03_0_03X | 614336252 | 92303 | Negative Control 1 | 847922 | 210013071 | 6.8 | 2025-10-26T05:53:05+0100 | Done | MTB NOT DETECTED | 18.8 |  | NEG | 3 | 0 | 03 | 03X | 0 | FALSE |  |  |  |  | Observation | INFO | 145
Xpert MTB-RIF Ultra | 4 | NEG_03_0_04X | 614336830 | 92303 | Negative Control 1 | 847922 | 657194 | 6.8 | 2025-10-26T05:53:30+0100 | Done | MTB NOT DETECTED | 18.7 |  | NEG | 3 | 0 | 04 | 04X | 0 | FALSE |  |  |  |  | Observation | INFO | 146
Xpert MTB-RIF Ultra | 4 | NEG_03_0_05X | 614336556 | 92303 | Negative Control 1 | 847922 | 702829 | 6.8 | 2025-10-26T05:53:53+0100 | Done | MTB NOT DETECTED | 17.3 |  | NEG | 3 | 0 | 05 | 05X | 0 | FALSE |  |  |  |  | Observation | INFO | 147
Xpert MTB-RIF Ultra | 4 | NEG_03_0_06+ | 614336421 | 92303 | Negative Control 1 | 847922 | 750444 | 6.8 | 2025-10-26T05:54:16+0100 | Done | MTB Trace DETECTED;RIF Resistance INDETERMINATE | 13.1 |  | NEG | 3 | 0 | 06 | 06+ | 0 | FALSE |  |  |  |  | Observation | INFO | 148
Xpert MTB-RIF Ultra | 4 | NEG_03_0_07+ | 614335759 | 92303 | Negative Control 1 | 847922 | 220000161 | 6.8 | 2025-10-26T05:54:41+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | 16.6 |  | NEG | 3 | 0 | 07 | 07+ | 0 | FALSE |  |  |  |  | Observation | INFO | 149
Xpert MTB-RIF Ultra | 4 | NEG_03_0_08+ | 614336979 | 92303 | Negative Control 1 | 847922 | 760491 | 6.8 | 2025-10-26T05:55:05+0100 | Done | MTB NOT DETECTED | 17.9 |  | NEG | 3 | 0 | 08 | 08+ | 0 | FALSE |  |  |  |  | Observation | INFO | 150
Xpert MTB-RIF Ultra | 4 | NEG_03_0_09+ | 614336667 | 92303 | Negative Control 1 | 847922 | 210068154 | 6.8 | 2025-10-26T05:55:30+0100 | Done | MTB NOT DETECTED | 17.7 |  | NEG | 3 | 0 | 09 | 09+ | 0 | FALSE |  |  |  |  | Observation | INFO | 151
Xpert MTB-RIF Ultra | 4 | NEG_03_0_10+ | 614336051 | 92303 | Negative Control 1 | 847922 | 902140 | 6.8 | 2025-10-26T05:55:53+0100 | Done | MTB NOT DETECTED | 16.2 |  | NEG | 3 | 0 | 10 | 10+ | 0 | FALSE |  |  |  |  | Observation | INFO | 152
Xpert MTB-RIF Ultra | 4 | NEG_04_0_01X | 614340140 | 92303 | Negative Control 1 | 847922 | 210097747 | 6.8 | 2025-10-26T09:34:16+0100 | Done | MTB NOT DETECTED | 19.4 |  | NEG | 4 | 0 | 01 | 01X | 0 | FALSE |  |  |  |  | Observation | INFO | 153
Xpert MTB-RIF Ultra | 4 | NEG_04_0_02X | 614339872 | 92303 | Negative Control 1 | 847922 | 210013071 | 6.8 | 2025-10-26T09:34:40+0100 | Done | MTB NOT DETECTED | 17 |  | NEG | 4 | 0 | 02 | 02X | 0 | FALSE |  |  |  |  | Observation | INFO | 154
Xpert MTB-RIF Ultra | 4 | NEG_04_0_03X | 614339510 | 92303 | Negative Control 1 | 847922 | 657194 | 6.8 | 2025-10-26T09:35:05+0100 | Done | MTB NOT DETECTED | 18.2 |  | NEG | 4 | 0 | 03 | 03X | 0 | FALSE |  |  |  |  | Observation | INFO | 155
Xpert MTB-RIF Ultra | 4 | NEG_04_0_04X | 614340440 | 92303 | Negative Control 1 | 847922 | 702829 | 6.8 | 2025-10-26T09:35:30+0100 | Done | MTB NOT DETECTED | 15.9 |  | NEG | 4 | 0 | 04 | 04X | 0 | FALSE |  |  |  |  | Observation | INFO | 156
Xpert MTB-RIF Ultra | 4 | NEG_04_0_05X | 614339216 | 92303 | Negative Control 1 | 847922 | 220000161 | 6.8 | 2025-10-26T09:35:54+0100 | Done | MTB NOT DETECTED | 16.8 |  | NEG | 4 | 0 | 05 | 05X | 0 | FALSE |  |  |  |  | Observation | INFO | 157
Xpert MTB-RIF Ultra | 4 | NEG_04_0_06+ | 614339977 | 92303 | Negative Control 1 | 847922 | 760491 | 6.8 | 2025-10-26T09:36:20+0100 | Done | MTB NOT DETECTED | 19.2 |  | NEG | 4 | 0 | 06 | 06+ | 0 | FALSE |  |  |  |  | Observation | INFO | 158
Xpert MTB-RIF Ultra | 4 | NEG_04_0_07+ | 614340317 | 92303 | Negative Control 1 | 847922 | 210068154 | 6.8 | 2025-10-26T09:36:44+0100 | Done | MTB NOT DETECTED | 17.6 |  | NEG | 4 | 0 | 07 | 07+ | 0 | FALSE |  |  |  |  | Observation | INFO | 159
Xpert MTB-RIF Ultra | 4 | NEG_04_0_08+ | 614339723 | 92303 | Negative Control 1 | 847922 | 701892 | 6.8 | 2025-10-26T09:37:09+0100 | Done | MTB NOT DETECTED | 17.1 |  | NEG | 4 | 0 | 08 | 08+ | 0 | FALSE |  |  |  |  | Observation | INFO | 160
Xpert MTB-RIF Ultra | 4 | NEG_04_0_09+ | 614339343 | 92303 | Negative Control 1 | 847922 | 866331 | 6.8 | 2025-10-26T09:37:32+0100 | Done | MTB NOT DETECTED | 17.1 |  | NEG | 4 | 0 | 09 | 09+ | 0 | FALSE |  |  |  |  | Observation | INFO | 161
Xpert MTB-RIF Ultra | 4 | NEG_04_0_10+ | 614338967 | 92303 | Positive Control 1 | 847922 | 902140 | 6.8 | 2025-10-26T09:37:56+0100 | Done | MTB NOT DETECTED | 20.5 |  | NEG | 4 | 0 | 10 | 10+ | 0 | FALSE |  |  |  |  | Observation | INFO | 162
Xpert MTB-RIF Ultra | 4 | NEG_05_0_01X | 614343144 | 92303 | Negative Control 1 | 847922 | 775845 | 6.8 | 2025-10-26T09:42:20+0100 | Done | MTB NOT DETECTED | 21.1 |  | NEG | 5 | 0 | 01 | 01X | 0 | FALSE |  |  |  |  | Observation | INFO | 163
Xpert MTB-RIF Ultra | 4 | NEG_05_0_02X | 614343860 | 92303 | Negative Control 1 | 847922 | 667585 | 6.8 | 2025-10-26T09:42:44+0100 | Done | MTB NOT DETECTED | 20.1 |  | NEG | 5 | 0 | 02 | 02X | 0 | FALSE |  |  |  |  | Observation | INFO | 164
Xpert MTB-RIF Ultra | 4 | NEG_05_0_03X | 614343496 | 92303 | Negative Control 1 | 847922 | 665815 | 6.8 | 2025-10-26T09:43:10+0100 | Done | MTB NOT DETECTED | 18.9 |  | NEG | 5 | 0 | 03 | 03X | 0 | FALSE |  |  |  |  | Observation | INFO | 165
Xpert MTB-RIF Ultra | 4 | NEG_05_0_04X | 614342592 | 92303 | Negative Control 1 | 847922 | 633201 | 6.8 | 2025-10-26T09:43:34+0100 | Done | MTB NOT DETECTED | 21.8 |  | NEG | 5 | 0 | 04 | 04X | 0 | FALSE |  |  |  |  | Observation | INFO | 166
Xpert MTB-RIF Ultra | 4 | NEG_05_0_05X | 614342822 | 92303 | Negative Control 1 | 847922 | 865593 | 6.8 | 2025-10-26T09:43:58+0100 | Done | MTB NOT DETECTED | 22.3 |  | NEG | 5 | 0 | 05 | 05X | 0 | FALSE |  |  |  |  | Observation | INFO | 167
Xpert MTB-RIF Ultra | 4 | NEG_05_0_06+ | 614343315 | 92303 | Negative Control 1 | 847922 | 791277 | 6.8 | 2025-10-26T09:44:22+0100 | Done | MTB NOT DETECTED | 25.8 |  | NEG | 5 | 0 | 06 | 06+ | 0 | FALSE |  |  |  |  | Observation | INFO | 168
Xpert MTB-RIF Ultra | 4 | NEG_05_0_07+ | 614343631 | 92303 | Negative Control 1 | 847922 | 636287 | 6.8 | 2025-10-26T09:44:46+0100 | Done | MTB NOT DETECTED | 25.7 |  | NEG | 5 | 0 | 07 | 07+ | 0 | FALSE |  |  |  |  | Observation | INFO | 169
Xpert MTB-RIF Ultra | 4 | NEG_05_0_08+ | 614342981 | 92303 | Negative Control 1 | 847922 | 730163 | 6.8 | 2025-10-26T09:45:11+0100 | Done | MTB NOT DETECTED | 20 |  | NEG | 5 | 0 | 08 | 08+ | 0 | FALSE |  |  |  |  | Observation | INFO | 170
Xpert MTB-RIF Ultra | 4 | NEG_05_0_09+ | 614342683 | 92303 | Negative Control 1 | 847922 | 670395 | 6.8 | 2025-10-26T09:45:36+0100 | Done | MTB NOT DETECTED | 16.2 |  | NEG | 5 | 0 | 09 | 09+ | 0 | FALSE |  |  |  |  | Observation | INFO | 171
Xpert MTB-RIF Ultra | 4 | NEG_05_0_10+ | 614342383 | 92303 | Negative Control 1 | 847922 | 210047638 | 6.8 | 2025-10-26T09:46:00+0100 | Done | MTB NOT DETECTED | 17.9 |  | NEG | 5 | 0 | 10 | 10+ | 0 | FALSE |  |  |  |  | Observation | INFO | 172
Xpert MTB-RIF Ultra | 4 | NEG_06_0_01X | 614346890 | 92303 | Negative Control 1 | 847922 | 886194 | 6.8 | 2025-10-26T14:28:54+0100 | Done | MTB NOT DETECTED | 18.3 |  | NEG | 6 | 0 | 01 | 01X | 0 | FALSE |  |  |  |  | Observation | INFO | 173
Xpert MTB-RIF Ultra | 4 | NEG_06_0_02X | 614347344 | 92303 | Negative Control 1 | 847922 | 797004 | 6.8 | 2025-10-26T14:29:19+0100 | Done | MTB NOT DETECTED | 28.2 |  | NEG | 6 | 0 | 02 | 02X | 0 | FALSE |  |  |  |  | Observation | INFO | 174
Xpert MTB-RIF Ultra | 4 | NEG_06_0_03X | 614347612 | 92303 | Negative Control 1 | 847922 | 742313 | 6.8 | 2025-10-26T14:29:43+0100 | Done | MTB NOT DETECTED | 22.9 |  | NEG | 6 | 0 | 03 | 03X | 0 | FALSE |  |  |  |  | Observation | INFO | 175
Xpert MTB-RIF Ultra | 4 | NEG_06_0_04X | 614345900 | 92303 | Negative Control 1 | 847922 | 604077 | 6.8 | 2025-10-26T14:30:07+0100 | Done | MTB NOT DETECTED | 17.7 |  | NEG | 6 | 0 | 04 | 04X | 0 | FALSE |  |  |  |  | Observation | INFO | 176
Xpert MTB-RIF Ultra | 4 | NEG_06_0_05X | 614346594 | 92303 | Negative Control 1 | 847922 | 617583 | 6.8 | 2025-10-26T14:30:32+0100 | Done | MTB NOT DETECTED | 17.5 |  | NEG | 6 | 0 | 05 | 05X | 0 | FALSE |  |  |  |  | Observation | INFO | 177
Xpert MTB-RIF Ultra | 4 | NEG_06_0_06+ | 614347035 | 92303 | Negative Control 1 | 847922 | 725382 | 6.8 | 2025-10-26T14:30:56+0100 | Done | MTB NOT DETECTED | 28.3 |  | NEG | 6 | 0 | 06 | 06+ | 0 | FALSE |  |  |  |  | Observation | INFO | 178
Xpert MTB-RIF Ultra | 4 | NEG_06_0_07+ | 614346097 | 92303 | Negative Control 1 | 847922 | 710982 | 6.8 | 2025-10-26T14:31:22+0100 | Done | MTB NOT DETECTED | 19.7 |  | NEG | 6 | 0 | 07 | 07+ | 0 | FALSE |  |  |  |  | Observation | INFO | 179
Xpert MTB-RIF Ultra | 4 | NEG_06_0_08+ | 614347803 | 92303 | Negative Control 1 | 847922 | 657194 | 6.8 | 2025-10-26T14:31:46+0100 | Done | MTB NOT DETECTED | 24.1 |  | NEG | 6 | 0 | 08 | 08+ | 0 | FALSE |  |  |  |  | Observation | INFO | 180
Xpert MTB-RIF Ultra | 4 | NEG_06_0_09+ | 614346729 | 92303 | Negative Control 1 | 847922 | 702829 | 6.8 | 2025-10-26T14:32:09+0100 | Done | MTB NOT DETECTED | 18.8 |  | NEG | 6 | 0 | 09 | 09+ | 0 | FALSE |  |  |  |  | Observation | INFO | 181
Xpert MTB-RIF Ultra | 4 | NEG_06_0_10+ | 614347201 | 92303 | Negative Control 1 | 847922 | 220000161 | 6.8 | 2025-10-26T14:32:31+0100 | Done | MTB NOT DETECTED | 21.2 |  | NEG | 6 | 0 | 10 | 10+ | 0 | FALSE |  |  |  |  | Observation | INFO | 182
Xpert MTB-RIF Ultra | 4 | NEG_07_0_01X | 614351832 | 92303 | Negative Control 1 | 847922 | 886194 | 6.8 | 2025-10-26T15:58:21+0100 | Done | MTB NOT DETECTED | 29.5 |  | NEG | 7 | 0 | 01 | 01X | 0 | FALSE |  |  |  |  | Observation | INFO | 183
Xpert MTB-RIF Ultra | 4 | NEG_07_0_02X | 614351590 | 92303 | Negative Control 1 | 847922 | 797004 | 6.8 | 2025-10-26T15:58:46+0100 | Done | MTB NOT DETECTED | 27.6 |  | NEG | 7 | 0 | 02 | 02X | 0 | FALSE |  |  |  |  | Observation | INFO | 184
Xpert MTB-RIF Ultra | 4 | NEG_07_0_03X | 614351220 | 92303 | Negative Control 1 | 847922 | 742313 | 6.8 | 2025-10-26T15:59:10+0100 | Done | MTB NOT DETECTED | 25.6 |  | NEG | 7 | 0 | 03 | 03X | 0 | FALSE |  |  |  |  | Observation | INFO | 185
Xpert MTB-RIF Ultra | 4 | NEG_07_0_04X | 614350974 | 92303 | Negative Control 1 | 847922 | 604077 | 6.8 | 2025-10-26T15:59:32+0100 | Done | MTB NOT DETECTED | 23.7 |  | NEG | 7 | 0 | 04 | 04X | 0 | FALSE |  |  |  |  | Observation | INFO | 186
Xpert MTB-RIF Ultra | 4 | NEG_07_0_05X | 614350660 | 92303 | Negative Control 1 | 847922 | 617583 | 6.8 | 2025-10-26T15:59:57+0100 | Done | MTB NOT DETECTED | 23.9 |  | NEG | 7 | 0 | 05 | 05X | 0 | FALSE |  |  |  |  | Observation | INFO | 187
Xpert MTB-RIF Ultra | 4 | NEG_07_0_06+ | 614350499 | 92303 | Negative Control 1 | 847922 | 725382 | 6.8 | 2025-10-26T16:00:22+0100 | Done | MTB NOT DETECTED | 24.4 |  | NEG | 7 | 0 | 06 | 06+ | 0 | FALSE |  |  |  |  | Observation | INFO | 188
Xpert MTB-RIF Ultra | 4 | NEG_07_0_07+ | 614351129 | 92303 | Negative Control 1 | 847922 | 710982 | 6.8 | 2025-10-26T16:00:48+0100 | Done | MTB NOT DETECTED | 27 |  | NEG | 7 | 0 | 07 | 07+ | 0 | FALSE |  |  |  |  | Observation | INFO | 189
Xpert MTB-RIF Ultra | 4 | NEG_07_0_08+ | 614350803 | 92303 | Negative Control 1 | 847922 | 657194 | 6.8 | 2025-10-26T16:01:10+0100 | Done | MTB NOT DETECTED | 24 |  | NEG | 7 | 0 | 08 | 08+ | 0 | FALSE |  |  |  |  | Observation | INFO | 190
Xpert MTB-RIF Ultra | 4 | NEG_07_0_09+ | 614351713 | 92303 | Negative Control 1 | 847922 | 702829 | 6.8 | 2025-10-26T16:01:34+0100 | Done | MTB NOT DETECTED | 25.4 |  | NEG | 7 | 0 | 09 | 09+ | 0 | FALSE |  |  |  |  | Observation | INFO | 191
Xpert MTB-RIF Ultra | 4 | NEG_07_0_10+ | 614351397 | 92303 | Negative Control 1 | 847922 | 220000161 | 6.8 | 2025-10-26T16:01:56+0100 | Done | MTB NOT DETECTED | 23.4 |  | NEG | 7 | 0 | 10 | 10+ | 0 | FALSE |  |  |  |  | Observation | INFO | 192
Xpert MTB-RIF Ultra | 4 | NEG_08_0_01X | 614383328 | 92303 | Negative Control 1 | 847922 | 886194 | 6.8 | 2025-10-26T18:28:41+0100 | Done | MTB NOT DETECTED | 28.9 |  | NEG | 8 | 0 | 01 | 01X | 0 | FALSE |  |  |  |  | Observation | INFO | 193
Xpert MTB-RIF Ultra | 4 | NEG_08_0_02X | 614383866 | 92303 | Negative Control 1 | 847922 | 797004 | 6.8 | 2025-10-26T18:29:31+0100 | Done | MTB NOT DETECTED | 16.6 |  | NEG | 8 | 0 | 02 | 02X | 0 | FALSE |  |  |  |  | Observation | INFO | 194
Xpert MTB-RIF Ultra | 4 | NEG_08_0_03X | 614352842 | 92303 | Negative Control 1 | 847922 | 742313 | 6.8 | 2025-10-26T18:30:44+0100 | Done | MTB NOT DETECTED | 21.6 |  | NEG | 8 | 0 | 03 | 03X | 0 | FALSE |  |  |  |  | Observation | INFO | 195
Xpert MTB-RIF Ultra | 4 | NEG_08_0_04X | 614384298 | 92303 | Negative Control 1 | 847922 | 604077 | 6.8 | 2025-10-26T18:31:55+0100 | Done | MTB NOT DETECTED | 23.1 |  | NEG | 8 | 0 | 04 | 04X | 0 | FALSE |  |  |  |  | Observation | INFO | 196
Xpert MTB-RIF Ultra | 4 | NEG_08_0_05X | 614352910 | 92303 | Negative Control 1 | 847922 | 617583 | 6.8 | 2025-10-26T18:33:10+0100 | Done | MTB NOT DETECTED | 19.4 |  | NEG | 8 | 0 | 05 | 05X | 0 | FALSE |  |  |  |  | Observation | INFO | 197
Xpert MTB-RIF Ultra | 4 | NEG_08_0_06+ | 614383509 | 92303 | Negative Control 1 | 847922 | 725382 | 6.8 | 2025-10-26T18:33:34+0100 | Done | MTB NOT DETECTED | 25.2 |  | NEG | 8 | 0 | 06 | 06+ | 0 | FALSE |  |  |  |  | Observation | INFO | 198
Xpert MTB-RIF Ultra | 4 | NEG_08_0_07+ | 614382571 | 92303 | Negative Control 1 | 847922 | 710982 | 6.8 | 2025-10-26T18:34:00+0100 | Done | MTB NOT DETECTED | 28.7 |  | NEG | 8 | 0 | 07 | 07+ | 0 | FALSE |  |  |  |  | Observation | INFO | 199
Xpert MTB-RIF Ultra | 4 | NEG_08_0_08+ | 614383987 | 92303 | Negative Control 1 | 847922 | 912152 | 6.8 | 2025-10-26T18:34:23+0100 | Done | MTB NOT DETECTED | 16.7 |  | NEG | 8 | 0 | 08 | 08+ | 0 | FALSE |  |  |  |  | Observation | INFO | 200
Xpert MTB-RIF Ultra | 4 | NEG_08_0_09+ | 614384411 | 92303 | Negative Control 1 | 847922 | 784559 | 6.8 | 2025-10-26T18:34:48+0100 | Done | MTB NOT DETECTED | 17.4 |  | NEG | 8 | 0 | 09 | 09+ | 0 | FALSE |  |  |  |  | Observation | INFO | 201
Xpert MTB-RIF Ultra | 4 | NEG_08_0_10+ | 614352459 | 92303 | Negative Control 1 | 847922 | 210053180 | 6.8 | 2025-10-26T18:35:13+0100 | Done | MTB NOT DETECTED | 29.2 |  | NEG | 8 | 0 | 10 | 10+ | 0 | FALSE |  |  |  |  | Observation | INFO | 202
Xpert MTB-RIF Ultra | 4 | NEG_09_0_01X | 614387708 | 92303 | Negative Control 1 | 847922 | 686063 | 6.8 | 2025-10-26T21:41:21+0100 | Done | MTB NOT DETECTED | 18.1 |  | NEG | 9 | 0 | 01 | 01X | 0 | FALSE |  |  |  |  | Observation | INFO | 203
Xpert MTB-RIF Ultra | 4 | NEG_09_0_02X | 614388006 | 92303 | Negative Control 1 | 847922 | 684250 | 6.8 | 2025-10-26T21:41:45+0100 | Done | MTB NOT DETECTED | 19.8 |  | NEG | 9 | 0 | 02 | 02X | 0 | FALSE |  |  |  |  | Observation | INFO | 204
Xpert MTB-RIF Ultra | 4 | NEG_09_0_03X | 614387446 | 92303 | Negative Control 1 | 847922 | 790976 | 6.8 | 2025-10-26T21:42:10+0100 | Done | MTB NOT DETECTED | 16.6 |  | NEG | 9 | 0 | 03 | 03X | 0 | FALSE |  |  |  |  | Observation | INFO | 205
Xpert MTB-RIF Ultra | 4 | NEG_09_0_04X | 614387172 | 92303 | Negative Control 1 | 847922 | 654187 | 6.8 | 2025-10-26T21:42:34+0100 | Done | MTB NOT DETECTED | 19.2 |  | NEG | 9 | 0 | 04 | 04X | 0 | FALSE |  |  |  |  | Observation | INFO | 206
Xpert MTB-RIF Ultra | 4 | NEG_09_0_05X | 614386812 | 92303 | Negative Control 1 | 847922 | 674666 | 6.8 | 2025-10-26T21:42:57+0100 | Done | MTB NOT DETECTED | 19 |  | NEG | 9 | 0 | 05 | 05X | 0 | FALSE |  |  |  |  | Observation | INFO | 207
Xpert MTB-RIF Ultra | 4 | NEG_09_0_06+ | 614386569 | 92303 | Negative Control 1 | 847922 | 797456 | 6.8 | 2025-10-26T21:43:20+0100 | Done | MTB NOT DETECTED | 18.1 |  | NEG | 9 | 0 | 06 | 06+ | 0 | FALSE |  |  |  |  | Observation | INFO | 208
Xpert MTB-RIF Ultra | 4 | NEG_09_0_07+ | 614387311 | 92303 | Negative Control 1 | 847922 | 695878 | 6.8 | 2025-10-26T21:43:46+0100 | Done | MTB NOT DETECTED | 21.8 |  | NEG | 9 | 0 | 07 | 07+ | 0 | FALSE |  |  |  |  | Observation | INFO | 209
Xpert MTB-RIF Ultra | 4 | NEG_09_0_08+ | 614387595 | 92303 | Negative Control 1 | 847922 | 784143 | 6.8 | 2025-10-26T21:44:12+0100 | Done | MTB NOT DETECTED | 18.9 |  | NEG | 9 | 0 | 08 | 08+ | 0 | FALSE |  |  |  |  | Observation | INFO | 210
Xpert MTB-RIF Ultra | 4 | NEG_09_0_09+ | 614387921 | 92303 | Negative Control 1 | 847922 | 673187 | 6.8 | 2025-10-26T21:44:35+0100 | Done | MTB NOT DETECTED | 21.9 |  | NEG | 9 | 0 | 09 | 09+ | 0 | FALSE |  |  |  |  | Observation | INFO | 211
Xpert MTB-RIF Ultra | 4 | NEG_09_0_10+ | 614386955 | 92303 | Negative Control 1 | 847922 | 888731 | 6.8 | 2025-10-26T21:44:59+0100 | Done | MTB NOT DETECTED | 18.5 |  | NEG | 9 | 0 | 10 | 10+ | 0 | FALSE |  |  |  |  | Observation | INFO | 212
Xpert MTB-RIF Ultra | 4 | NEG_10_0_01X | 614390504 | 92303 | Negative Control 1 | 847922 | 657194 | 6.8 | 2025-10-26T23:37:12+0100 | Done | MTB NOT DETECTED | 20.5 |  | NEG | 10 | 0 | 01 | 01X | 0 | FALSE |  |  |  |  | Observation | INFO | 213
Xpert MTB-RIF Ultra | 4 | NEG_10_0_02X | 614390490 | 92303 | Negative Control 1 | 847922 | 702829 | 6.8 | 2025-10-26T23:38:00+0100 | Done | MTB NOT DETECTED | 16.4 |  | NEG | 10 | 0 | 02 | 02X | 0 | FALSE |  |  |  |  | Observation | INFO | 214
Xpert MTB-RIF Ultra | 4 | NEG_10_0_03X | 614390034 | 92303 | Negative Control 1 | 847922 | 220000161 | 6.8 | 2025-10-26T23:39:12+0100 | Done | MTB NOT DETECTED | 20.9 |  | NEG | 10 | 0 | 03 | 03X | 0 | FALSE |  |  |  |  | Observation | INFO | 215
Xpert MTB-RIF Ultra | 4 | NEG_10_0_04X | 614390494 | 92303 | Negative Control 1 | 847922 | 912152 | 6.8 | 2025-10-26T23:40:24+0100 | Done | MTB NOT DETECTED | 17.1 |  | NEG | 10 | 0 | 04 | 04X | 0 | FALSE |  |  |  |  | Observation | INFO | 216
Xpert MTB-RIF Ultra | 4 | NEG_10_0_05X | 614390402 | 92303 | Negative Control 1 | 847922 | 210046047 | 6.8 | 2025-10-26T23:41:37+0100 | Done | MTB NOT DETECTED | 20.5 |  | NEG | 10 | 0 | 05 | 05X | 0 | FALSE |  |  |  |  | Observation | INFO | 217
Xpert MTB-RIF Ultra | 4 | NEG_10_0_06+ | 614390505 | 92303 | Negative Control 1 | 847922 | 701892 | 6.8 | 2025-10-26T23:42:50+0100 | Aborted | ERROR | 92 | Error 2008: Syringe pressure reading of 100.1 PSI exceeds the protocol limit of 100.0 PSI | NEG | 10 | 0 | 06 | 06+ | 0 | FALSE |  |  | 2008 | MINOR_FUNCTIONAL_ERROR | Minor Functional Error | WARN | 218
Xpert MTB-RIF Ultra | 4 | NEG_10_0_07+ | 614390503 | 92303 | Negative Control 1 | 847922 | 866331 | 6.8 | 2025-10-26T23:43:14+0100 | Done | MTB NOT DETECTED | 13.4 |  | NEG | 10 | 0 | 07 | 07+ | 0 | FALSE |  |  |  |  | Observation | INFO | 219
Xpert MTB-RIF Ultra | 4 | NEG_10_0_08+ | 614389891 | 92303 | Negative Control 1 | 847922 | 902140 | 6.8 | 2025-10-26T23:43:37+0100 | Done | MTB NOT DETECTED | 17.1 |  | NEG | 10 | 0 | 08 | 08+ | 0 | FALSE |  |  |  |  | Observation | INFO | 220
Xpert MTB-RIF Ultra | 4 | NEG_10_0_09+ | 614390529 | 92303 | Negative Control 1 | 847922 | 635023 | 6.8 | 2025-10-26T23:44:01+0100 | Done | MTB NOT DETECTED | 18.1 |  | NEG | 10 | 0 | 09 | 09+ | 0 | FALSE |  |  |  |  | Observation | INFO | 221
Xpert MTB-RIF Ultra | 4 | NEG_10_0_10+ | 614390273 | 92303 | Negative Control 1 | 847922 | 652001 | 6.8 | 2025-10-26T23:44:25+0100 | Done | MTB NOT DETECTED | 18.2 |  | NEG | 10 | 0 | 10 | 10+ | 0 | FALSE |  |  |  |  | Observation | INFO | 222

## Seal Test Info
Seal Test Information
| Neg | Pos
ROBAL | ROBAL-006 | ROBAL-006
Part Number | 700-5702 | 700-5702
Batch Number | 1001506040 | 1001506046 | Mismatch
Cartridge Number (LSP) | 92302 | 92303 | Mismatch
PO Number | 1110580 | 1110580
Assay Family | MTB | MTB |  | Utrustningsfilen
Weight Loss Spec | 100 | 100
Balance ID Number | SCLSYS01, 003070 | SCLSYS03, SCLSYS04 |  | SCLSYS01, 003070 | SCLSYS03, SCLSYS04
Balance Cal Due Date | Apr-26 | Apr-26 |  | Apr-26 | Apr-26
Vacuum Oven ID Number | 003478, 003298 | 003346, 003587 |  | 003478, 003298 | 003346, 003587
Vacuum Oven Cal Due Date | Jun-26 | Jun-26 |  | Jun-26 | Jun-26
Timer ID Number | Nr. 18, Nr. 19 | Nr. 20, Nr. 21 |  | Nr. 18, Nr. 19 | Nr. 20, Nr. 21
Timer Cal Due Date | Mar-26 | Mar-26 |  | Mar-26 | Mar-26
Name of Tester
| Bincenth Salazar | Bincenth Salazar
| David Björklund | Katarina Ilic
| Evangelos Samakidis | Laith Ukra
| Mahdi Hussein | Leila Hussein
| Marina Ahmadi | Marina Ahmadi
| Warda Yousaf | Toni Shabo
Print Full Name, Sign, and Date
| Neg | Pos
| Test Testsson, JESP, 2026 06 06 | Test Testsson, JESP, 2026 06 06

## STF Sum
| Datasheet | Cartridge ID | Initial Weight (mg) | Final Weight (mg) | Weight Loss (mg) | Weight Loss Pass/Fail | Observation
NEG | Datasheet (1) | 01_0_08 | 37055.7 | 33197.2 | 3858.5 | FAIL | Lid Weld (Leak)
NEG | Datasheet (4) | 07_0_03 | 37645.1 | 37650 | -4.9 | Minusvärde | NA
POS | Datasheet (1) | 02_2_20 | 37472.4 | 37476.1 | -3.7 | Minusvärde | NA
POS | Datasheet (2) | 04_1_14 | 37668.5 | 37518.5 | 150 | FAIL | Lid Weld (Leak)
POS | Datasheet (3) | 06_1_13 | 37480.5 | 36479.9 | 1000.6 | FAIL | NA

## Infinity GX
Utrustningsfilen
NEG | [7 empty] | Pos
INSTRUMENT | [7 empty] | INSTRUMENT
GeneXpert / Infinity | GX5 | GX6 | Infinity-VI | Infinity-VIII |  |  |  | GeneXpert / Infinity | GX1 | GX2 | GX3 | GX7 | Infinity-I | Infinity-III | Infinity-V
Calibration Due Date | 2026-04-25 | 2026-04-25 | 2026-05-25 | 2025-01-26 |  |  |  | Calibration Due Date | Mar-26 | Mar-26 | Mar-26 | Mar-26 | Jun-26 | Apr-26 | 2026-01-01
PIPETTER | [7 empty] | PIPETTER
Pipette Information | Nr. 6 | Nr. 7 | Nr. 23 | Nr. 32 | Nr. 33 | Nr. 35 | Nr. 36 | Pipette Information | Nr. 3 | Nr. 4 | Nr. 13 | Nr. 17 | Nr. 18 | Nr. 19 | Nr. 20
Calibration Due Date | 2026-02-25 | 2026-02-25 | 2026-03-25 | 2026-01-01 | 2026-01-01 | 2026-02-25 | 2026-02-25 | Calibration Due Date | 2026-01-01 | 2025-12-25 | 2026-03-25 | 2026-01-25 | 2026-03-25 | 2026-03-25 | 2026-03-25
[8 empty] | Pipette Information | Nr. 21 | Nr. 22 | Nr. 24 | Nr. 25 | Nr. 26 | Nr. 34 | Nr. 37
[8 empty] | Calibration Due Date | 2026-03-25 | 2026-01-25 | 2026-03-25 | 2026-03-25 | 2026-03-25 | 2026-02-01 | Sep-25
Uppdaterad: 2025 12 08 / JESP

## MTB ULTRA
MTB ULTRA
P/N | Artikel | Lotnr. | Utgångsdatum | Förvaringsplats | Antal i lager | Senast inventering | SIGN
001-1923 | INTROL TB1081MDR2 | B27FEB25A | 2027-02-28 | IPT kylskåp | 5 Lådor | 2025-11-28 | JESP
001-1923 |  | N/A | N/A |  | N/A | N/A | N/A
001-1923 |  | N/A | N/A |  | N/A | N/A | N/A
001-1923 |  | N/A | N/A |  | N/A | N/A | N/A
P/N | Artikel | Lotnr. | Utgångsdatum | Förvaringsplats | Antal i lager | Senast inventering | SIGN
001-1968 | INTROL TBWT-04 MTB Ultra | B01NOV24B | 2026-11-30 | IPT kylskåp | 5.5 lådor | 2025-11-24 | ELS
001-1968 |  | N/A | N/A |  | N/A | N/A | N/A
001-1968 |  | N/A | N/A |  | N/A | N/A | N/A
001-1968 |  | N/A | N/A |  | N/A | N/A | N/A
MTB - TET ¦ 500-0995
P/N 500-0995 MTB TET Buffer (exp 6 mån) 500-0995, BUFFER, TET, 50_0.1_0.1
Lotnummer |  |  | Utgångsdatum |  | Volym, L | Senast inventering | SIGN
1001510545 |  |  | 2026-05-14 |  | 25L | 2025-11-21 | ELS
1001503562 |  |  | 2026-03-30 |  | 10L | 2025-11-24 | vdao
N/A |  |  | N/A |  | N/A | N/A | N/A
N/A |  |  | N/A |  | N/A | N/A | N/A
MTB - SR ¦ 500-2608
P/N 500-2608 MTB SR (exp 4 mån) 500-2608, IPT BULK REAGENT, SAMPLE TREATMENT, MTB
Lotnummer |  |  | Utgångsdatum |  | Volym, L | Senast inventering | SIGN
1001508989 |  |  | 2026-03-06 |  | 50L | 2025-11-21 | ELS
1001505209 |  |  | 2026-02-10 |  | 30 | 2025-11-18 | JESP
N/A |  |  | N/A |  | N/A | N/A | N/A
N/A |  |  | N/A |  | N/A | N/A | N/A

## SharePoint Info
SharePoint Information
Work Center | ROBAL 06
Order# | 1110580
SAP Batch# | 1001506046
LSP | 92303
Material | GXMTB/RIF-ULTRA-50
BBD/SLED | 2027-04-11
ROBAL - Actual start date/time | 2025-10-25 22:40
Sample Reagent use | 1001501891, 1001502674
Sample Reagent P/N | 700-8870
Order quantity | 29000
ROBAL - Till Packning | 27958
IPT - Testing Finalized | 2025-10-26 23:00
MES Order | JA