## Intended Use
Intended Use
This worksheet is associated with MTB Ultra V2 Product QC Testing Procedure (D25862).
Fields in blue must be input electronically.
Fields in green may be entered manually.
Fields in yellow perform calculations.
Mean Ct Evaluation
This tab calculates the mean Ct for SPC and the mean IS1081-IS6110 Ct from Negative samples. It will calculate the mean Ct and mean TM for IS1081-IS6110, rpo1, rpo2, rpo4 from Sample Type 1 as well as the mean Ct and mean TM for IS1081-IS6110, rpo1, rpo2, rpo3, rpo4 from Sample Type 2 and compares them to specifications on "Specifications" tab. Failing results are highlighted in RED.
This tab performs the same calculations for the "First 10" results | [7 empty] | Table of Tabs to Print-SV only
Vacuum Seal Data | [7 empty] | Tab | Print or Not Print
This tab is used to record and evaluate the Vacuum Seal Test data from ROBAL in-process testing. | [7 empty] | Intended Use | Do NOT print
| Column A | [6 empty] | Mean Ct Evaluation | Print
| Enumerates the sampled cartridges. | [6 empty] | Vacuum Seal Data | Print
[8 empty] | Test Summary | Print
| Column B | [6 empty] | Data Summary | Print
| Records the Cartridge ID. | [6 empty] | Data- First 10 | Print
[8 empty] | Data- Sample Type 0 | Print
| Column C | [6 empty] | Data- Sample Type 1 | Print
| Records the results of seal test. | [6 empty] | Data- Sample Type 2 | Print
| Valid entries are: | [6 empty] | Data- Replaced | Print
|  | "P" | [5 empty] | Statistical Process Control | Print
|  | "F" | [5 empty] | Resample Vacuum Seal Data | Print if applicable
|  | "N" | [5 empty] | Resample Mean Ct Evaluation | Print if applicable
[8 empty] | Resample Data Summary | Print if applicable
| Column D | [6 empty] | Data- Resample Replaced | Print if applicable
| Calculates the results of seal test. | [6 empty] | Specifications | Do NOT print
| The result is "NA" if the test replicate was not tested or does not apply. | [6 empty] | HistoricalAccuStats | Do NOT print
| The results of the calculations are used as a counter for internal worksheet purposes. | [6 empty] | For Reference Only | Do NOT print
Test Summary
B3 | Select Catalog Number from pull down menu
B4 | Enter lot number (without leading 0)
Row 9 and 11 | Enter Control lot number
Row 15-16 | Enter pipet information
Row 18-21 | Enter GeneXpert information
P18 | Enter bags tested using infinity
A23-G32 | Table 4
| column C | Counts individual error types from data summary tab
| column D | Compares number of failures to pass/resample criteria (N>4 resample)
| column E | Counts individual error types from resample data summary tab
| coulmn F | Compares number of failures to pass/resample criteria (N>10 resample)
C34 | Counts the max count for Single Minor Failure
C36 | Max allowable failure per n=210
C34 | Counts the max count for Single Minor Failure for Resample Data
C36 | Max allowable failure per n=410
I23-K26 | Table 5
K24 | Manually enter number of failures for Minor Visual Failure
K25 | Manually enter number of failures for Major Visual Failure
K26 | Manually enter number of failures for Delamination
I29-M33 | Table 6
L31 | Counts number of False Negs for original Data
M31 | Counts number of False Negs for Resample Data
L32 | Counts number of False Pos for original Data
M32 | Counts number of False Pos for Resample Data
L33 | Counts number of Major Visual Failures for original Data
M33 | Counts number of Major Visual Failures for Resample Data
Row 40-49 | Table 7
| Column D | Rows 41-42 | Counts number of replicates and valid replicates from Data Summary and Replaced Data tabs
|  | Row 43 | Counts total number of minor failures from data summary (pressure and minor visual failures) and table 4 (for other minors)
|  | Row 44 | Calculates total number of major failures from rows 35 and 36
|  | Row 45 | Calculates total number of failures from rows 34 and 37
| Column E | Acceptance Criteria per Specifications tab
| Column F | Rows 44-45 | Evaluates number of failures against limits on the specifications tab
| Column H (For Resample Data ) | Rows 41-42 | Counts number of replicates and valid replicates from Data Summary and Replaced Data tabs
|  | Row 43 | Counts total number of minor failures from data summary (pressure and minor visual failures) and table 4 (for other minors)
|  | Row 44 | Calculates total number of major failures from rows 35 and 36
|  | Row 45 | Calculates total number of failures from rows 34 and 37
| Column J | Row 44 | Calculates combined failures for original and resample
| Column K | Rows 44-45 | Evaluates number of failures against limits on the specifications tab
D48 | Calculates overall result using table 4, table 5 and table 6
H48 | Pull down menu for NCR number if applicable
H49 | Pull down menu for LFI number if applicable
Data Summary
Rows 2-4 | Lot information pulled from Test Summary and Specifications tabs
D6 | Counts number of replicates in Rows 9-18
E6 | Compares Number of replicates to Specifications
B7 | Compares the lot number from Row 3 with the Data - First 10 tab (all replicates must match)
D7 | Compares Assay name from C4 with the Data - First 10 tab (all replicates must match)
Rows 10-19 | Column A | Pulls Sample ID from Data - First 10 tab
| Column B | Pulls Test Result from Data - First 10 tab
| Column C | Evaluates the result in column B and Data - First 10 tab in the following order;
|  |  | Found |  | Report | Uses Conditional formatting
|  |  | Blank Cell |  | Blank | Green | Pass
|  |  | Error Code Present |  | Error Text | Yellow | Minor Failure or Data Entry Error
|  |  | Invalid |  | Invalid | Red | Major Failure
|  |  | If S |  | Seal Test Fail
|  |  | If V |  | Minor Visual
|  |  | If M |  | Major Visual
|  |  | if B |  | Barcode Fail
|  |  | if D |  | Delamination
|  |  | Text not recognized |  | Data Entry Error
|  |  | Correct Test Result |  | Pass
|  |  | Incorrect Test Result |  | Major Functional
| Column D | Evaluates maximum pressure from Data - First 10 tab versus limit in Specifications tab
Uses Conditional formatting
|  |  | Green | Pass
|  |  | Yellow | N/A
|  |  | Red | Fail
| Column E | Comment either with pull down menu or free text
D21 | Counts number of replicates in Rows 23-122
B22 | Compares the lot number from Row 3 with the Data - Sample Type 0 tab (all replicates must match)
D22 | Compares Assay name from C4 with the Data - Sample Type 0 tab (all replicates must match)
Rows 25-124 | Column A | Pulls Sample ID from Data - Sample Type 0 tab
| Column B | Pulls Test Result from Data - Sample Type 0 tab
| Column C | Evaluates the result in column B and Data - Sample Type 0 tab in the following order;
|  |  | Found |  | Report | Uses Conditional formatting
|  |  | Blank Cell |  | Blank | Green | Pass
|  |  | Error Code Present |  | Error Text | Yellow | Minor Failure or Data Entry Error
|  |  | Invalid |  | Invalid | Red | Major Failure
|  |  | If S |  | Seal Test Fail
|  |  | If V |  | Minor Visual
|  |  | If M |  | Major Visual
|  |  | if B |  | Barcode Fail
|  |  | if D |  | Delamination
|  |  | Text not recognized |  | Data Entry Error
|  |  | Correct Test Result |  | Pass
|  |  | Incorrect Test Result |  | Major Functional
| Column D | Evaluates maximum pressure from Data - Sample Type 0 tab versus limit in Specifications tab
Uses Conditional formatting
|  |  | Green | Pass
|  |  | Yellow | N/A
|  |  | Red | Fail
| Column E | Comment either with pull down menu or free text
D126 | Counts number of replicates in Rows 130-210
F124 | Compares Number of replicates to Specifications
B127 | Compares the lot number from Row 3 with the Data - Sample Type 1 tab (all replicates must match)
D127 | Compares Assay name from C4 with the Data - Sample Type 1 tab (all replicates must match)
Rows 130-209 | Column A | Pulls Sample ID from Data - Sample Type 1 tab
| Column B | Pulls Test Result from Data - Sample Type 1 tab
| Column C | Evaluates the result in column B and Data - Sample Type 1 tab in the following order;
|  |  | Found |  | Report | Uses Conditional formatting
|  |  | Blank Cell |  | Blank | Green | Pass
|  |  | Error Code Present |  | Error Text | Yellow | Minor Failure or Data Entry Error
|  |  | Invalid |  | Invalid | Red | Major Failure
|  |  | If V |  | Minor Visual
|  |  | If M |  | Major Visual
|  |  | if B |  | Barcode Fail
|  |  | if D |  | Delamination
|  |  | Text not recognized |  | Data Entry Error
|  |  | If SPC Ct=0 |  | SPC Ct=0
|  |  | Correct Test Result |  | Pass
|  |  | Incorrect Test Result |  | Major Functional
| Column D | Evaluates maximum pressure from Data - Sample Type 0 tab versus limit in Specifications tab
Uses Conditional formatting
|  |  | Green | Pass
|  |  | Yellow | N/A
|  |  | Red | Fail
| Column E | Comment either with pull down menu or free text
D211 | Counts number of replicates in Rows 215-234
B212 | Compares the lot number from Row 3 with the Data - Sample Type 1 tab (all replicates must match)
D212 | Compares Assay name from C4 with the Data - Sample Type 1 tab (all replicates must match)
Rows 215-234 | Column A | Pulls Sample ID from Data - Sample Type 1 tab
| Column B | Pulls Test Result from Data - Sample Type 1 tab
| Column C | Evaluates the result in column B and Data - Sample Type 1 tab in the following order;
|  |  | Found |  | Report | Uses Conditional formatting
|  |  | Blank Cell |  | Blank | Green | Pass
|  |  | Error Code Present |  | Error Text | Yellow | Minor Failure or Data Entry Error
|  |  | Invalid |  | Invalid | Red | Major Failure
|  |  | If V |  | Minor Visual
|  |  | If M |  | Major Visual
|  |  | if B |  | Barcode Fail
|  |  | if D |  | Delamination
|  |  | Text not recognized |  | Data Entry Error
|  |  | If SPC Ct=0 |  | SPC Ct=0
|  |  | Correct Test Result |  | Pass
|  |  | Incorrect Test Result |  | Major Functional
| Column D | Evaluates maximum pressure from Data - Sample Type 0 tab versus limit in Specifications tab
Uses Conditional formatting
|  |  | Green | Pass
|  |  | Yellow | N/A
|  |  | Red | Fail
| Column E | Comment either with pull down menu or free text
D236 | Counts number of replaced replicates in Data - Replaced tab
row 238-337 | Pull data from Data - Replaced Tab as necessary
| Column E | Comment either with pull down menu or free text
Data - First 10
This tab provides a location to record the GX test summary results for First 10 Cartridges.
Paste in data to Row 2 to Row 11
No calculations are performed in this tab.
Data - Sample Type 0
This tab provides a location to record the GX test summary results for Sample Type 0.
Paste in data starting with Row 2
No calculations are performed in this tab.
Paste resample data starting from row 104
Data - Sample Type 1
This tab provides a location to record the GX test summary results for Sample Type 1.
Paste in data starting with Row 2
No calculations are performed in this tab.
Paste resample data starting from row 84
Data - Sample Type 2
This tab provides a location to record the GX test summary results for Sample Type 2.
Paste in data starting with Row 2
No calculations are performed in this tab.
Paste resample data starting from row 24
Data - Replaced
This tab provides a location to record the GX test summary results for replicates that were replaced.
Paste in data starting with Row 2
No calculations are performed in this tab.
Statistical Process Control
This tab perfoms intra- and inter-sample point statistical data analyses.
The data are evaluated relative to acceptable limits established through historical data with a specified confidence.
The confidence becomes more preceise as more data is accumulated during the course of the run.
The Blue Data Line indicates the actual process.
The Red Control Line indicates the historical confidence level.
The Red Control Line is calculated using the Default Sample Size but will adjust based on the actual sample size.
After a Sample Size of Subgroup is entered, this value is used to calculate the remainder of the Red Control Line.
The process is operating at or better than historical limits as long as the Blue Data Line remains below the Red Control Line. | Statistical Justification
The equation for the Red Control Line is based on Exact Binomial Confidence Limits.
The Red Highlight (of the entered Default Sample Size) is based on a normal approximation.
Justification for these equations can be found in: Hollander, Myles and Wolfe, Douglas, Nonparametric Statsitical Methods, Page 23 to 24, John Wilely and Sons, NY, NY 1973.
The Excel equations can be found in the HistoricalAccurStats tab.
These results can be duplicated in Minitab version 15, using STATS/Basic Statistics/1-Proportion.
Resample Mean Ct Evaluation
This tab calculates the mean Ct for SPC and the mean IS1081-IS6110 Ct from Negative samples. It will calculate the mean Ct and mean TM for IS1081-IS6110, rpo1, rpo2, rpo4 from Sample Type 1 as well as the mean Ct and mean TM for IS1081-IS6110, rpo1, rpo2, rpo3, rpo4 from Sample Type 2 and compares them to specifications on "Specifications" tab. Failing results are highlighted in RED.
This tab performs the same calculations for the "First 10" results
Resample Vacuum Seal Data
This tab is used to record and evaluate the Vacuum Seal Test data from ROBAL in-process testing.
Column A
Enumerates the sampled cartridges.
Column B
Records the Cartridge ID.
Column C
Records the results of seal test.
Valid entries are:
"P"
"F"
"N"
Column D
Calculates the results of seal test.
The result is "NA" if the test replicate was not tested or does not apply.
The results of the calculations are used as a counter for internal worksheet purposes.
Resample Data Summary
Rows 2-4 | Lot information pulled from Test Summary and Specifications tabs
D6 | Counts number of replicates in Rows 9-18
E6 | Compares Number of replicates to Specifications
B7 | Compares the lot number from Row 3 with the Data - First 10 tab (all replicates must match)
D7 | Compares Assay name from C4 with the Data - First 10 tab (all replicates must match)
Rows 10-19 | Column A | Pulls Sample ID from Data - First 10 tab
| Column B | Pulls Test Result from Data - First 10 tab
| Column C | Evaluates the result in column B and Data - First 10 tab in the following order;
|  |  | Found |  | Report | Uses Conditional formatting
|  |  | Blank Cell |  | Blank | Green | Pass
|  |  | Error Code Present |  | Error Text | Yellow | Minor Failure or Data Entry Error
|  |  | Invalid |  | Invalid | Red | Major Failure
|  |  | If S |  | Seal Test Fail
|  |  | If V |  | Minor Visual
|  |  | If M |  | Major Visual
|  |  | if B |  | Barcode Fail
|  |  | if D |  | Delamination
|  |  | Text not recognized |  | Data Entry Error
|  |  | Correct Test Result |  | Pass
|  |  | Incorrect Test Result |  | Major Functional
| Column D | Evaluates maximum pressure from Data - First 10 tab versus limit in Specifications tab
Uses Conditional formatting
|  |  | Green | Pass
|  |  | Yellow | N/A
|  |  | Red | Fail
| Column E | Comment either with pull down menu or free text
D21 | Counts number of replicates in Rows 23-122
B22 | Compares the lot number from Row 3 with the Data - Sample Type 0 tab (all replicates must match)
D22 | Compares Assay name from C4 with the Data - Sample Type 0 tab (all replicates must match)
Rows 25-124 | Column A | Pulls Sample ID from Data - Sample Type 0 tab
| Column B | Pulls Test Result from Data - Sample Type 0 tab
| Column C | Evaluates the result in column B and Data - Sample Type 0 tab in the following order;
|  |  | Found |  | Report | Uses Conditional formatting
|  |  | Blank Cell |  | Blank | Green | Pass
|  |  | Error Code Present |  | Error Text | Yellow | Minor Failure or Data Entry Error
|  |  | Invalid |  | Invalid | Red | Major Failure
|  |  | If S |  | Seal Test Fail
|  |  | If V |  | Minor Visual
|  |  | If M |  | Major Visual
|  |  | if B |  | Barcode Fail
|  |  | if D |  | Delamination
|  |  | Text not recognized |  | Data Entry Error
|  |  | Correct Test Result |  | Pass
|  |  | Incorrect Test Result |  | Major Functional
| Column D | Evaluates maximum pressure from Data - Sample Type 0 tab versus limit in Specifications tab
Uses Conditional formatting
|  |  | Green | Pass
|  |  | Yellow | N/A
|  |  | Red | Fail
| Column E | Comment either with pull down menu or free text
D126 | Counts number of replicates in Rows 130-210
F124 | Compares Number of replicates to Specifications
B127 | Compares the lot number from Row 3 with the Data - Sample Type 1 tab (all replicates must match)
D127 | Compares Assay name from C4 with the Data - Sample Type 1 tab (all replicates must match)
Rows 130-209 | Column A | Pulls Sample ID from Data - Sample Type 1 tab
| Column B | Pulls Test Result from Data - Sample Type 1 tab
| Column C | Evaluates the result in column B and Data - Sample Type 1 tab in the following order;
|  |  | Found |  | Report | Uses Conditional formatting
|  |  | Blank Cell |  | Blank | Green | Pass
|  |  | Error Code Present |  | Error Text | Yellow | Minor Failure or Data Entry Error
|  |  | Invalid |  | Invalid | Red | Major Failure
|  |  | If V |  | Minor Visual
|  |  | If M |  | Major Visual
|  |  | if B |  | Barcode Fail
|  |  | if D |  | Delamination
|  |  | Text not recognized |  | Data Entry Error
|  |  | If SPC Ct=0 |  | SPC Ct=0
|  |  | Correct Test Result |  | Pass
|  |  | Incorrect Test Result |  | Major Functional
| Column D | Evaluates maximum pressure from Data - Sample Type 0 tab versus limit in Specifications tab
Uses Conditional formatting
|  |  | Green | Pass
|  |  | Yellow | N/A
|  |  | Red | Fail
| Column E | Comment either with pull down menu or free text
D211 | Counts number of replicates in Rows 215-234
B212 | Compares the lot number from Row 3 with the Data - Sample Type 1 tab (all replicates must match)
D212 | Compares Assay name from C4 with the Data - Sample Type 1 tab (all replicates must match)
Rows 215-234 | Column A | Pulls Sample ID from Data - Sample Type 1 tab
| Column B | Pulls Test Result from Data - Sample Type 1 tab
| Column C | Evaluates the result in column B and Data - Sample Type 1 tab in the following order;
|  |  | Found |  | Report | Uses Conditional formatting
|  |  | Blank Cell |  | Blank | Green | Pass
|  |  | Error Code Present |  | Error Text | Yellow | Minor Failure or Data Entry Error
|  |  | Invalid |  | Invalid | Red | Major Failure
|  |  | If V |  | Minor Visual
|  |  | If M |  | Major Visual
|  |  | if B |  | Barcode Fail
|  |  | if D |  | Delamination
|  |  | Text not recognized |  | Data Entry Error
|  |  | If SPC Ct=0 |  | SPC Ct=0
|  |  | Correct Test Result |  | Pass
|  |  | Incorrect Test Result |  | Major Functional
| Column D | Evaluates maximum pressure from Data - Sample Type 0 tab versus limit in Specifications tab
Uses Conditional formatting
|  |  | Green | Pass
|  |  | Yellow | N/A
|  |  | Red | Fail
| Column E | Comment either with pull down menu or free text
D236 | Counts number of replaced replicates in Data - Replaced tab
row 238-337 | Pull data from Data - Replaced Tab as necessary
| Column E | Comment either with pull down menu or free text
Data -Resample Replaced
This tab provides a location to record the GX test summary results for replicates that were replaced.
Paste in data starting with Row 2
No calculations are performed in this tab.
Specifications
This tab lists the specifications that are used to analyze the data input in the Cartridge Data and Results Summary tabs.
No calculations are performed in this tab.
HistoricalAccuStats
This tab supports the calculations performed in the Statistical Process Control tab.
For Reference Only
This tab informs pull-down menus in other tabs.
No calculations are performed in this tab.
Document Revision History
Use this table format and only provide the changes applicable to this revision and only the one prior. (Main purpose of listing the current changes is to aid in the training via Compliance Wire, provide high level list of changes, help with auditing)
Rev. | Effective: | Description of Changes
AW | Current | Added P/N 700-9436 to Specification tab & revised worksheet formulas as needed.
AV | 2025-07-11 | No changes to this document

## Test Summary
Product QC Testing Summary
Product Catalog/Part No.: | GXMTBRIFULTRA50 |  |  | (Select using drop-down menu.) |  |  | Type of Resample |  | NA
Cartridge No. (LSP) | 92303 |  |  | Expiration Date |  | 2027-04-11
Table 1; Test and Control Material Information
TB1081-MDR2 Cells |  | TET Buffer |  | Sample Reagent |  |  | Sample Reagent
Part No.* | Lot No. | Part No.* | Lot No. | Part No.* | Lot No. | Bag # | Part No.* | Lot No. | Bag #
001-1923 | C15NOV23D | 500-1315 or 500-0995 | 1001500034 | 500-2095, 500-2290, 500-2608, 700-5208, 700-6052 or 700-9287,700-8870 | 1001499559 | 00+05-10+TBWT | 500-2095, 500-2290, 500-2608, 700-5208, 700-6052 or 700-9287,700-8870 | 1001502674 | 01-02
TBWT Cells |  |  |  | Sample Reagent |  |  | Sample Reagent
Part No.* | Lot No. |  |  | Part No.* | Lot No. | Bag # | Part No.* | Lot No. | Bag #
001-1968 | B01NOV24D |  |  | 500-2095, 500-2290, 500-2608, 700-5208, 700-6052 or 700-9287,700-8870 | 1001501891 | 03-04 | 500-2095, 500-2290, 500-2608, 700-5208, 700-6052 or 700-9287,700-8870 | N/A | N/A
* Circle one, if more than 1 option is displayed.
Table 2; Pipette Information
Cepheid ID | Nr. 35 |  | Nr. 21 |  | Nr. 22 |  | N/A |  | N/A |  | N/A
Calibration Due Date | 2025-11-25 |  | 2025-12-25 |  | 2026-01-25 |  | N/A |  | N/A |  | N/A
Table 3; GeneXperts Used:
Cepheid ID | Infinity-I |  | Infinity-V |  | Infinity-VIII |  | Infinity-VI |  | N/A |  | N/A |  | Bag #s Tested on Inifinity | N/A
Calibration Due Date | Dec-25 |  | 2026-01-01 |  | 2026-01-01 |  | 2025-11-25 |  | N/A |  | N/A
Cepheid ID | N/A |  | N/A |  | N/A |  | N/A |  | N/A |  | N/A
Calibration Due Date | N/A |  | N/A |  | N/A |  | N/A |  | N/A |  | N/A
Table 4: Single Minor Failure Mode: | [7 empty] | Table 5: Visual and Delamination Resample Results
Type |  | Original Count | Result | Resample Count | Final Count | Result |  | Type |  | Resample Count
Invalid (w/o error code) and SPC Ct=0 |  | 1 | Pass | NA | NA | NA |  | Minor Visual Failure |  | N/A
Curve Fit Error |  | 0 | Pass | NA | NA | NA |  | Major Visual Failure |  | N/A
Probe Check Low |  | 1 | Pass | NA | NA | NA |  | Delamination |  | N/A
Probe Check High |  | 1 | Pass | NA | NA | NA |  | Note: Enter in results if applicable (all resamples other than functional and full)
CIT and SLD |  | 1 | Pass | NA | NA | NA
Pressure Abort or Max Pressure > 90 PSI |  | 1 | Pass | NA | NA | NA |  | Table 6: Major Failure Results
Sample Volume Adequacy |  | 0 | Pass | NA | NA | NA |  | Type |  |  | Original Count | Resample Count
Delamination |  | 0 | Pass | NA | NA | NA |  | No. False Negative |  |  | 0 | NA
[8 empty] | No. False Positive |  |  | 1 | NA
[8 empty] | No. Major Visual Inspection Failures: |  |  | 0 | NA
| Max Count | 1 |  |  | NA |  |  | MTB Trace Detected |  |  | 1 | NA
| Max Acceptable: | 4 |  |  | 10
| (Out of ) | (210) |  |  | (410)
Table 7: Result Summary:
Type |  |  | Original Count | Accept Criteria | Result |  | Resample Count | Accept Criteria | Combined Count | Result
Total No. Cartridges Registered: |  |  | 217 | NA | NA |  | NA | NA | NA | NA
Total No. Valid Test Replicates: |  |  | 210 | NA | NA |  | NA | NA | NA | NA
No. Minor Failures: |  |  | 7 | NA | NA |  | NA | NA | NA | NA
No. Seal Test Failures: |  |  | 4 | ≤9 | Pass |  | NA | NA | NA | NA
No. Total Major Failures: |  |  | 2 | ≤ 1 Resample at 2-3 | Resample |  | NA | ≤ 5 | NA | NA
No. Total Failures: |  |  | 9 | ≤ 9 | Pass |  | NA | ≤ 10 | NA | NA
|  | Original Result: | Resample per D12547 |  |  | NC# | N/A
|  | Overall Result: | Choose correct resample type |  |  | LFI# | N/A
| Comments | [9 empty] | ¨N/A
Loten har även testats av Evangelos Samakidis.
| Recorded By: | Jesper Fredriksson | [5 empty] | Date: | 2025 10 26
| PQC Reviewed By: | Fredrik Jespersson | [5 empty] | Date: | 2025 10 27

## Vacuum Seal Data
| Cartridge ID | P/F/N | 0/1/NA
1 | 01_0_01 | P | 0
2 | 01_0_02 | P | 0
3 | 01_0_03 | P | 0
4 | 01_0_04 | P | 0
5 | 01_0_05 | P | 0
6 | 01_0_06 | P | 0
7 | 01_0_07 | P | 0
8 | 01_0_08 | F | 1
9 | 01_0_09 | P | 0
10 | 01_0_10 | P | 0
11 | 01_1_11 | P | 0
12 | 01_1_12 | P | 0
13 | 01_1_13 | P | 0
14 | 01_1_14 | P | 0
15 | 01_1_15 | P | 0
16 | 01_1_16 | P | 0
17 | 01_1_17 | P | 0
18 | 01_1_18 | P | 0
19 | 01_2_19 | P | 0
20 | 01_2_20 | P | 0
21 | 02_0_01 | P | 0
22 | 02_0_02 | P | 0
23 | 02_0_03 | P | 0
24 | 02_0_04 | P | 0
25 | 02_0_05 | P | 0
26 | 02_0_06 | P | 0
27 | 02_0_07 | P | 0
28 | 02_0_08 | P | 0
29 | 02_0_09 | P | 0
30 | 02_0_10 | P | 0
31 | 02_1_11 | P | 0
32 | 02_1_12 | P | 0
33 | 02_1_13 | P | 0
34 | 02_1_14 | P | 0
35 | 02_1_15 | P | 0
36 | 02_1_16 | P | 0
37 | 02_1_17 | P | 0
38 | 02_1_18 | P | 0
39 | 02_2_19 | P | 0
40 | 02_2_20 | P | 0
41 | 03_0_01 | P | 0
42 | 03_0_02 | P | 0
43 | 03_0_03 | P | 0
44 | 03_0_04 | P | 0
45 | 03_0_05 | P | 0
46 | 03_0_06 | P | 0
47 | 03_0_07 | P | 0
48 | 03_0_08 | P | 0
49 | 03_0_09 | P | 0
50 | 03_0_10 | P | 0
51 | 03_1_11 | P | 0
52 | 03_1_12 | P | 0
53 | 03_1_13 | P | 0
54 | 03_1_14 | P | 0
55 | 03_1_15 | P | 0
56 | 03_1_16 | P | 0
57 | 03_1_17 | P | 0
58 | 03_1_18 | P | 0
59 | 03_2_19 | P | 0
60 | 03_2_20 | P | 0
61 | 04_0_01 | P | 0
62 | 04_0_02 | P | 0
63 | 04_0_03 | P | 0
64 | 04_0_04 | P | 0
65 | 04_0_05 | P | 0
66 | 04_0_06 | P | 0
67 | 04_0_07 | P | 0
68 | 04_0_08 | P | 0
69 | 04_0_09 | P | 0
70 | 04_0_10 | P | 0
71 | 04_1_11 | P | 0
72 | 04_1_12 | P | 0
73 | 04_1_13 | P | 0
74 | 04_1_14 | F | 1
75 | 04_1_15 | P | 0
76 | 04_1_16 | P | 0
77 | 04_1_17 | P | 0
78 | 04_1_18 | P | 0
79 | 04_2_19 | P | 0
80 | 04_2_20 | P | 0
81 | 05_0_01 | P | 0
82 | 05_0_02 | P | 0
83 | 05_0_03 | P | 0
84 | 05_0_04 | P | 0
85 | 05_0_05 | P | 0
86 | 05_0_06 | P | 0
87 | 05_0_07 | P | 0
88 | 05_0_08 | P | 0
89 | 05_0_09 | P | 0
90 | 05_0_10 | P | 0
91 | 05_1_11 | P | 0
92 | 05_1_12 | P | 0
93 | 05_1_13 | P | 0
94 | 05_1_14 | P | 0
95 | 05_1_15 | P | 0
96 | 05_1_16 | P | 0
97 | 05_1_17 | P | 0
98 | 05_1_18 | P | 0
99 | 05_2_19 | P | 0
100 | 05_2_20 | P | 0
101 | 06_0_01 | P | 0
102 | 06_0_02 | P | 0
103 | 06_0_03 | P | 0
104 | 06_0_04 | P | 0
105 | 06_0_05 | P | 0
106 | 06_0_06 | P | 0
107 | 06_0_07 | P | 0
108 | 06_0_08 | P | 0
109 | 06_0_09 | P | 0
110 | 06_0_10 | P | 0
111 | 06_1_11 | P | 0
112 | 06_1_12 | P | 0
113 | 06_1_13 | F | 1
114 | 06_1_14 | P | 0
115 | 06_1_15 | P | 0
116 | 06_1_16 | F | 1
117 | 06_1_17 | P | 0
118 | 06_1_18 | P | 0
119 | 06_2_19 | P | 0
120 | 06_2_20 | P | 0
121 | 07_0_01 | P | 0
122 | 07_0_02 | P | 0
123 | 07_0_03 | P | 0
124 | 07_0_04 | P | 0
125 | 07_0_05 | P | 0
126 | 07_0_06 | P | 0
127 | 07_0_07 | P | 0
128 | 07_0_08 | P | 0
129 | 07_0_09 | P | 0
130 | 07_0_10 | P | 0
131 | 07_1_11 | P | 0
132 | 07_1_12 | P | 0
133 | 07_1_13 | P | 0
134 | 07_1_14 | P | 0
135 | 07_1_15 | P | 0
136 | 07_1_16 | P | 0
137 | 07_1_17 | P | 0
138 | 07_1_18 | P | 0
139 | 07_2_19 | P | 0
140 | 07_2_20 | P | 0
141 | 08_0_01 | P | 0
142 | 08_0_02 | P | 0
143 | 08_0_03 | P | 0
144 | 08_0_04 | P | 0
145 | 08_0_05 | P | 0
146 | 08_0_06 | P | 0
147 | 08_0_07 | P | 0
148 | 08_0_08 | P | 0
149 | 08_0_09 | P | 0
150 | 08_0_10 | P | 0
151 | 08_1_11 | P | 0
152 | 08_1_12 | P | 0
153 | 08_1_13 | P | 0
154 | 08_1_14 | P | 0
155 | 08_1_15 | P | 0
156 | 08_1_16 | P | 0
157 | 08_1_17 | P | 0
158 | 08_1_18 | P | 0
159 | 08_2_19 | P | 0
160 | 08_2_20 | P | 0
161 | 09_0_01 | P | 0
162 | 09_0_02 | P | 0
163 | 09_0_03 | P | 0
164 | 09_0_04 | P | 0
165 | 09_0_05 | P | 0
166 | 09_0_06 | P | 0
167 | 09_0_07 | P | 0
168 | 09_0_08 | P | 0
169 | 09_0_09 | P | 0
170 | 09_0_10 | P | 0
171 | 09_1_11 | P | 0
172 | 09_1_12 | P | 0
173 | 09_1_13 | P | 0
174 | 09_1_14 | P | 0
175 | 09_1_15 | P | 0
176 | 09_1_16 | P | 0
177 | 09_1_17 | P | 0
178 | 09_1_18 | P | 0
179 | 09_2_19 | P | 0
180 | 09_2_20 | P | 0
181 | 10_0_01 | P | 0
182 | 10_0_02 | P | 0
183 | 10_0_03 | P | 0
184 | 10_0_04 | P | 0
185 | 10_0_05 | P | 0
186 | 10_0_06 | P | 0
187 | 10_0_07 | P | 0
188 | 10_0_08 | P | 0
189 | 10_0_09 | P | 0
190 | 10_0_10 | P | 0
191 | 10_1_11 | P | 0
192 | 10_1_12 | P | 0
193 | 10_1_13 | P | 0
194 | 10_1_14 | P | 0
195 | 10_1_15 | P | 0
196 | 10_1_16 | P | 0
197 | 10_1_17 | P | 0
198 | 10_1_18 | P | 0
199 | 10_2_19 | P | 0
200 | 10_2_20 | P | 0
201 | N/A | N | N/A
202 | N/A | N | N/A
203 | N/A | N | N/A
204 | N/A | N | N/A
205 | N/A | N | N/A
206 | N/A | N | N/A
207 | N/A | N | N/A
208 | N/A | N | N/A
209 | N/A | N | N/A
210 | N/A | N | N/A
211 | N/A | N | N/A
212 | N/A | N | N/A
213 | N/A | N | N/A
214 | N/A | N | N/A
215 | N/A | N | N/A
216 | N/A | N | N/A
217 | N/A | N | N/A
218 | N/A | N | N/A
219 | N/A | N | N/A
220 | N/A | N | N/A
221 | N/A | N | N/A
222 | N/A | N | N/A
223 | N/A | N | N/A
224 | N/A | N | N/A
225 | N/A | N | N/A
226 | N/A | N | N/A
227 | N/A | N | N/A
228 | N/A | N | N/A
229 | N/A | N | N/A
230 | N/A | N | N/A
231 | N/A | N | N/A
232 | N/A | N | N/A
233 | N/A | N | N/A
234 | N/A | N | N/A
235 | N/A | N | N/A
236 | N/A | N | N/A
237 | N/A | N | N/A
238 | N/A | N | N/A
239 | N/A | N | N/A
240 | N/A | N | N/A
241 | N/A | N | N/A
242 | N/A | N | N/A
243 | N/A | N | N/A
244 | N/A | N | N/A
245 | N/A | N | N/A
246 | N/A | N | N/A
247 | N/A | N | N/A
248 | N/A | N | N/A
249 | N/A | N | N/A
250 | N/A | N | N/A
| Recorded By: | Jesper Fredriksson |  | Date: | 2025 10 26
Comments:
N/A

## Data Summary
Data Summary
Part No. |  | GXMTBRIFULTRA50
Expected Lot Number |  | 92303
Expected Assay Name |  | Xpert MTB-RIF Ultra
First 10 Sample Type | 0 | No. Of Replicates | 10
Lot Check | VALID | Assay Name Check | VALID
No. Of Replicates Check | PASS | Expected Result | MTB NOT DETECTED
Cartridge ID | Result | Failure check | Pressure Check | Comment (drop-down menu available)
NEG_00_0_01X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_00_0_02X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_00_0_03X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_00_0_04X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_00_0_05X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_00_0_06+ | MTB NOT DETECTED | PASS | PASS | N/A
NEG_00_0_07+ | MTB NOT DETECTED | PASS | PASS | N/A
NEG_00_0_08+ | MTB NOT DETECTED | PASS | PASS | N/A
NEG_00_0_09+ | MTB NOT DETECTED | PASS | PASS | N/A
NEG_00_0_10+ | MTB NOT DETECTED | PASS | PASS | N/A
Sample Type | 0 | No. Of Replicates | 100
Lot Check | VALID | Assay Name Check | VALID
No. Of Replicates Check | PASS | Expected Result | MTB NOT DETECTED
Cartridge ID | Result | Failure check | Pressure Check | Comment (drop-down menu available)
NEG_01_0_01X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_01_0_02X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_01_0_03X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_01_0_04X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_01_0_05X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_01_0_06+ | MTB NOT DETECTED | PASS | PASS | N/A
NEG_01_0_07+ | MTB NOT DETECTED | PASS | PASS | N/A
NEG_01_0_08A+ | MTB NOT DETECTED | PASS | PASS | See Replaced Data
NEG_01_0_09+ | MTB NOT DETECTED | PASS | PASS | N/A
NEG_01_0_10+ | MTB NOT DETECTED | PASS | PASS | N/A
NEG_02_0_01X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_02_0_02X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_02_0_03X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_02_0_04X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_02_0_05X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_02_0_06+ | MTB NOT DETECTED | PASS | PASS | N/A
NEG_02_0_07+ | MTB NOT DETECTED | PASS | PASS | N/A
NEG_02_0_08+ | MTB NOT DETECTED | PASS | PASS | N/A
NEG_02_0_09+ | MTB NOT DETECTED | PASS | PASS | N/A
NEG_02_0_10+ | MTB NOT DETECTED | PASS | PASS | N/A
NEG_03_0_01X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_03_0_02X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_03_0_03X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_03_0_04X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_03_0_05X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_03_0_06+ | MTB Trace DETECTED;RIF Resistance INDETERMINATE | MAJOR FUNCTIONAL | PASS | N/A
NEG_03_0_07+ | MTB DETECTED LOW;RIF Resistance DETECTED | MAJOR FUNCTIONAL | PASS | N/A
NEG_03_0_08+ | MTB NOT DETECTED | PASS | PASS | N/A
NEG_03_0_09+ | MTB NOT DETECTED | PASS | PASS | N/A
NEG_03_0_10+ | MTB NOT DETECTED | PASS | PASS | N/A
NEG_04_0_01X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_04_0_02X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_04_0_03X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_04_0_04X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_04_0_05X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_04_0_06+ | MTB NOT DETECTED | PASS | PASS | N/A
NEG_04_0_07+ | MTB NOT DETECTED | PASS | PASS | N/A
NEG_04_0_08+ | MTB NOT DETECTED | PASS | PASS | N/A
NEG_04_0_09+ | MTB NOT DETECTED | PASS | PASS | N/A
NEG_04_0_10+ | MTB NOT DETECTED | PASS | PASS | N/A
NEG_05_0_01X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_05_0_02X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_05_0_03X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_05_0_04X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_05_0_05X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_05_0_06+ | MTB NOT DETECTED | PASS | PASS | N/A
NEG_05_0_07+ | MTB NOT DETECTED | PASS | PASS | N/A
NEG_05_0_08+ | MTB NOT DETECTED | PASS | PASS | N/A
NEG_05_0_09+ | MTB NOT DETECTED | PASS | PASS | N/A
NEG_05_0_10+ | MTB NOT DETECTED | PASS | PASS | N/A
NEG_06_0_01X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_06_0_02X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_06_0_03X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_06_0_04X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_06_0_05X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_06_0_06+ | MTB NOT DETECTED | PASS | PASS | N/A
NEG_06_0_07+ | MTB NOT DETECTED | PASS | PASS | N/A
NEG_06_0_08+ | MTB NOT DETECTED | PASS | PASS | N/A
NEG_06_0_09+ | MTB NOT DETECTED | PASS | PASS | N/A
NEG_06_0_10+ | MTB NOT DETECTED | PASS | PASS | N/A
NEG_07_0_01X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_07_0_02X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_07_0_03X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_07_0_04X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_07_0_05X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_07_0_06+ | MTB NOT DETECTED | PASS | PASS | N/A
NEG_07_0_07+ | MTB NOT DETECTED | PASS | PASS | N/A
NEG_07_0_08+ | MTB NOT DETECTED | PASS | PASS | N/A
NEG_07_0_09+ | MTB NOT DETECTED | PASS | PASS | N/A
NEG_07_0_10+ | MTB NOT DETECTED | PASS | PASS | N/A
NEG_08_0_01X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_08_0_02X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_08_0_03X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_08_0_04X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_08_0_05X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_08_0_06+ | MTB NOT DETECTED | PASS | PASS | N/A
NEG_08_0_07+ | MTB NOT DETECTED | PASS | PASS | N/A
NEG_08_0_08+ | MTB NOT DETECTED | PASS | PASS | N/A
NEG_08_0_09+ | MTB NOT DETECTED | PASS | PASS | N/A
NEG_08_0_10+ | MTB NOT DETECTED | PASS | PASS | N/A
NEG_09_0_01X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_09_0_02X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_09_0_03X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_09_0_04X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_09_0_05X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_09_0_06+ | MTB NOT DETECTED | PASS | PASS | N/A
NEG_09_0_07+ | MTB NOT DETECTED | PASS | PASS | N/A
NEG_09_0_08+ | MTB NOT DETECTED | PASS | PASS | N/A
NEG_09_0_09+ | MTB NOT DETECTED | PASS | PASS | N/A
NEG_09_0_10+ | MTB NOT DETECTED | PASS | PASS | N/A
NEG_10_0_01X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_10_0_02X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_10_0_03X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_10_0_04X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_10_0_05X | MTB NOT DETECTED | PASS | PASS | N/A
NEG_10_0_06+ | ERROR | Error 2008: Syringe pressure reading of 100.1 PSI exceeds the protocol limit of 100.0 PSI | FAIL | N/A
NEG_10_0_07+ | MTB NOT DETECTED | PASS | PASS | N/A
NEG_10_0_08+ | MTB NOT DETECTED | PASS | PASS | N/A
NEG_10_0_09+ | MTB NOT DETECTED | PASS | PASS | N/A
NEG_10_0_10+ | MTB NOT DETECTED | PASS | PASS | N/A
Sample Type | 1 | No. Of Replicates | 80
Lot Check | VALID | Assay Name Check | VALID
No. Of Replicates Check | PASS | Expected Result | MTB DETECTED *;Rif Resistance DETECTED
Cartridge ID | Result | Failure check | Pressure Check | Comment (drop-down menu available)
C15NOV23D_01_1_11X | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_01_1_12X | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_01_1_13X | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_01_1_14X | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_01_1_15+ | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_01_1_16+ | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_01_1_17+ | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_01_1_18+ | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_02_1_11X | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_02_1_12X | ERROR | Error 5007: [CT1] probe check failed. Probe check value of 17 for reading number 2 was below the minimum of 110¦Error 5007: [SPC] probe check failed. Probe check value of 19 for reading number 2 was below the minimum of 119 | PASS | N/A
C15NOV23D_02_1_13X | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_02_1_14X | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_02_1_15+ | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_02_1_16+ | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_02_1_17+ | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_02_1_18+ | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_03_1_11X | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_03_1_12X | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_03_1_13X | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_03_1_14AX | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | See Replaced Data
C15NOV23D_03_1_15A+ | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | See Replaced Data
C15NOV23D_03_1_16A+ | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | See Replaced Data
C15NOV23D_03_1_17+ | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_03_1_18+ | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_04_1_11X | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_04_1_12X | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_04_1_13X | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_04_1_14AX | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | See Replaced Data
C15NOV23D_04_1_15+ | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_04_1_16+ | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_04_1_17+ | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_04_1_18+ | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_05_1_11X | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_05_1_12X | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_05_1_13X | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_05_1_14X | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_05_1_15+ | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_05_1_16+ | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_05_1_17+ | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_05_1_18+ | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_06_1_11X | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_06_1_12X | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_06_1_13AAX | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | See Replaced Data
C15NOV23D_06_1_14X | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_06_1_15+ | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_06_1_16A+_D06 | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | See Replaced Data
C15NOV23D_06_1_17+ | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_06_1_18+ | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_07_1_11X | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_07_1_12X | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_07_1_13X_D2B,11,05 | ERROR | Error 5006: [QC-2] probe check failed. Probe check value of 247 for reading number 1 was above the maximum of 223¦Error 5006: [QC-2] probe check failed. Probe check value of 247 for reading number 3 was above the maximum of 223 | PASS | N/A
C15NOV23D_07_1_14X | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_07_1_15+ | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_07_1_16+ | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_07_1_17+ | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_07_1_18+ | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_08_1_11X | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_08_1_12X | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_08_1_13X | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_08_1_14X_D2B | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_08_1_15+ | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_08_1_16+_D05 | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_08_1_17+ | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_08_1_18+ | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_09_1_11X | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_09_1_12X | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_09_1_13X | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_09_1_14X | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_09_1_15+ | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_09_1_16+ | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_09_1_17+ | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_09_1_18+ | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_10_1_11X | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_10_1_12X | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_10_1_13X | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_10_1_14X | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_10_1_15+ | MTB DETECTED LOW;RIF Resistance INDETERMINATE | INDETERMINATE | PASS | N/A
C15NOV23D_10_1_16+ | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_10_1_17+ | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
C15NOV23D_10_1_18+ | MTB DETECTED LOW;RIF Resistance DETECTED | PASS | PASS | N/A
Sample Type | 2 | No. Of Replicates | 20
Lot Check | VALID | Assay Name Check | VALID
No. Of Replicates Check | PASS | Expected Result | MTB DETECTED *;Rif Resistance NOT DETECTED
Cartridge ID | Result | Failure check | Pressure Check | Comment (drop-down menu available)
B01NOV24B_01_2_19X | MTB DETECTED LOW;RIF Resistance NOT DETECTED | PASS | PASS | N/A
B01NOV24B_01_2_20+ | MTB DETECTED LOW;RIF Resistance NOT DETECTED | PASS | PASS | N/A
B01NOV24B_02_2_19X | MTB DETECTED LOW;RIF Resistance NOT DETECTED | PASS | PASS | N/A
B01NOV24B_02_2_20+ | MTB DETECTED LOW;RIF Resistance NOT DETECTED | PASS | PASS | N/A
B01NOV24B_03_2_19X | NO RESULT | Error 2037: The cartridge integrity test failed at valve position 1390. The pressure change of -0.6 PSI did not exceed the requirement of 4.0 PSI. The pressure increased from 3.5 PSI to 3.0 PSI during the test | PASS | N/A
B01NOV24B_03_2_20+ | MTB DETECTED LOW;RIF Resistance NOT DETECTED | PASS | PASS | N/A
B01NOV24B_04_2_19X | MTB DETECTED LOW;RIF Resistance NOT DETECTED | SPC CT=0 | PASS | N/A
B01NOV24B_04_2_20+ | MTB DETECTED LOW;RIF Resistance NOT DETECTED | PASS | PASS | N/A
B01NOV24B_05_2_19X | MTB DETECTED LOW;RIF Resistance NOT DETECTED | PASS | PASS | N/A
B01NOV24B_05_2_20+ | MTB DETECTED LOW;RIF Resistance NOT DETECTED | PASS | PASS | N/A
B01NOV24B_06_2_19X | MTB DETECTED LOW;RIF Resistance NOT DETECTED | PASS | PASS | N/A
06_2_20+ | V | MINOR VISUAL | N/A | Liquid under lid
B01NOV24B_07_2_19X | MTB DETECTED LOW;RIF Resistance NOT DETECTED | PASS | PASS | N/A
B01NOV24B_07_2_20+ | MTB DETECTED LOW;RIF Resistance NOT DETECTED | PASS | PASS | N/A
B01NOV24B_08_2_19X | MTB DETECTED LOW;RIF Resistance NOT DETECTED | PASS | PASS | N/A
B01NOV24B_08_2_20+ | MTB DETECTED LOW;RIF Resistance NOT DETECTED | PASS | PASS | N/A
B01NOV24B_09_2_19X | MTB DETECTED LOW;RIF Resistance NOT DETECTED | PASS | PASS | N/A
B01NOV24B_09_2_20+ | MTB DETECTED LOW;RIF Resistance NOT DETECTED | PASS | PASS | N/A
B01NOV24B_10_2_19X | MTB DETECTED LOW;RIF Resistance NOT DETECTED | PASS | PASS | N/A
B01NOV24B_10_2_20+ | MTB DETECTED LOW;RIF Resistance NOT DETECTED | PASS | PASS | N/A
Replaced | ¨N/A | No. Of Replicates | 7
Cartridge ID | Result | Failure check | Pressure Check | Comment (drop-down menu available)
01_0_08+ | S | Seal Test Failure | N/A | Seal Test Failure error replaced per D12547 Rev FK
04_1_14X | S | Seal Test Failure | N/A | Seal Test Failure error replaced per D12547 Rev FK
C15NOV23D_03_1_14X | ERROR | N/A | PASS | Instrument error replaced per D12547 Rev FK
C15NOV23D_03_1_15+ | ERROR | N/A | PASS | Instrument error replaced per D12547 Rev FK
C15NOV23D_03_1_16+ | ERROR | N/A | PASS | Instrument error replaced per D12547 Rev FK
C15NOV23D_06_1_13AX | ERROR | N/A | PASS | Instrument error replaced per D12547 Rev FK
C15NOV23D_06_1_13X | S | Seal Test Failure | N/A | Seal Test Failure error replaced per D12547 Rev FK
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
N/A | N/A | N/A | N/A | N/A
Recorded By: | Jesper Fredriksson | Date: | 2025 10 26

## Mean Ct Evaluation
[5 empty] | Sample Type 0 (Negative) |  |  |  |  | Sample Type 1 (TBMDR2) | [10 empty] | Sample Type 2 (TBWT)
Data Set | IS1081-IS6110 Ct | IS1081-IS6110 Ct Mean | SPC Ct | SPC Ct Mean | Data Set | IS1081-IS6110 Ct | IS1081-IS6110 Ct Mean per Data Set | SPC Ct | SPC mean per Data Set | Data Set | IS1081-IS6110 Ct | rpo1 Ct | rpo2 Ct | rpo4 Ct | rpo1 Tm | rpo1 Tm Mean | rpo2 Tm | rpo2 Tm Mean | rpo4 Tm | rpo4 Tm Mean | Data Set | IS1081-IS6110 Ct | rpo1 Ct | rpo2 Ct | rpo3 Ct | rpo4 Ct | rpo1 Tm | rpo1 Tm Mean | rpo2 Tm | rpo2 Tm Mean | rpo3 Tm | rpo3 Tm Mean | rpo4 Tm | rpo4 Tm Mean
First 10 | 40 | 40 | 24.3 | 24 | 1 | 40 | 40 | 26.3 | 24.7 | 1 | 27.2 | 25 | 25.9 | 23.5 | 69.3 | 69.2 | 70 | 69.9 | 73.5 | 73.5 | 1-5 | 22.4 | 24.7 | 25 | 25.8 | 27.2 | 69 | 69.1 | 0 | 0 | 75.5 | 75.6 | 66.8 | 66.9
| 40 |  | 24.3 |  |  | 40 |  | 24 |  |  | 25.6 | 23.5 | 24.8 | 22.3 | 69.2 |  | 70 |  | 73.5 |  |  | 21.1 | 23.6 | 23.6 | 24.7 | 25.8 | 69 |  | 0 |  | 75.5 |  | 66.9
| 40 |  | 24.5 |  |  | 40 |  | 23.9 |  |  | 26.3 | 23.5 | 24.6 | 22.2 | 69.1 |  | 69.9 |  | 73.4 |  |  | 23.1 | 25 | 25.1 | 26.4 | 27.4 | 69.3 |  | 0 |  | 75.8 |  | 67
| 40 |  | 23.2 |  |  | 40 |  | 23.7 |  |  | 26.4 | 23.8 | 24.9 | 22.6 | 69.1 |  | 69.8 |  | 73.3 |  |  | 23.8 | 26.1 | 26.3 | 27.8 | 28.7 | 69 |  | 0 |  | 75.4 |  | 66.6
| 40 |  | 21.6 |  |  | 40 |  | 27.1 |  |  | 26.8 | 24 | 24.7 | 22.6 | 69.2 | PASS | 69.8 | PASS | 73.4 | PASS
| 40 | PASS | 23.9 | PASS |  | 40 | PASS | 23.6 | PASS |  | 26.2 | 23.4 | 24.6 | 22.4 | 69.4 |  | 70.1 |  | 73.6 |  |  | 24.4 | 25.2 | 25.2 | 26.1 | 27.4 | 69.2 | PASS | 0 | PASS | 75.7 | PASS | 67 | PASS
| 40 |  | 24.5 |  |  | 40 |  | 23.3 |  |  | 26.8 | 24.6 | 26.1 | 23.2 | 69.3 |  | 69.9 |  | 73.5
| 40 |  | 24.6 |  |  | 40 |  | 25.8 |  |  | 25.9 | 23 | 24 | 21.6 | 69.3 |  | 69.9 |  | 73.5 |  |  | 23 | 25.2 | 25.1 | 26.1 | 26.6 | 69.1 |  | 0 |  | 75.6 |  | 66.8
| 40 |  | 24.4 |  |  | 40 |  | 25 |  | 2 | 25.8 | 23.8 | 23.8 | 22.1 | 69.3 | 69.3 | 70.1 | 70 | 73.6 | 73.5 |  | 22.7 | 24.3 | 24.5 | 25.8 | 26.7 | 69.2 |  | 0 |  | 75.7 |  | 67
| 40 |  | 25 |  |  | 40 |  | 24.2 | [13 empty] | 23 | 25.1 | 24.8 | 25.7 | 27.2 | 69.2 |  | 0 |  | 75.7 |  | 66.9
[5 empty] | 2 | 40 | 40 | 24.4 | 24.7 |  | 25.5 | 23.6 | 24.5 | 22.1 | 69.4 |  | 70.1 |  | 73.6 |  | 6-10 | 24 | 25.4 | 25.3 | 26.3 | 26.8 | 69.3 | 69.2 | 0 | 0 | 75.7 | 75.6 | 66.8 | 66.9
[5 empty] | 2 | 40 |  | 24.4 |  |  | 24.9 | 23 | 24.2 | 21.7 | 69.3 |  | 70.1 |  | 73.6
[6 empty] | 40 |  | 24.4 |  |  | 26.7 | 23.8 | 24.9 | 22.4 | 69.1 | PASS | 69.8 | PASS | 73.3 | PASS |  | 23 | 24.6 | 24.7 | 25.8 | 27.4 | 69.2 |  | 0 |  | 75.7 |  | 67
[6 empty] | 40 |  | 24.1 |  |  | 25.4 | 23 | 24.1 | 21.8 | 69.2 |  | 69.8 |  | 73.4 |  |  | 23.1 | 25.1 | 25 | 26.2 | 27.6 | 69.1 |  | 0 |  | 75.6 |  | 66.9
[6 empty] | 40 |  | 24.2 |  |  | 26.2 | 23.4 | 24.4 | 22.1 | 69.3 |  | 70 |  | 73.5 |  |  | 23.7 | 26.8 | 26.9 | 28.1 | 29.3 | 69.1 |  | 0 |  | 75.6 |  | 66.9
[6 empty] | 40 | PASS | 28.3 | PASS |  | 26.4 | 24.3 | 24.9 | 22.7 | 69.2 |  | 69.8 |  | 73.3 |  |  | 22.6 | 24.6 | 23.9 | 25.1 | 26.7 | 69.1 | PASS | 0 | PASS | 75.4 | PASS | 66.7 | PASS
[6 empty] | 40 |  | 25.3 |  | 3 | 26 | 23.5 | 24.4 | 21.8 | 69.4 | 69.4 | 70.1 | 70.1 | 73.6 | 73.6 |  | 21.9 | 24.1 | 23.9 | 25.1 | 25.8 | 69.1 |  | 0 |  | 75.5 |  | 66.9
[6 empty] | 40 |  | 24.5 |  |  | 26.8 | 24.6 | 25.6 | 23.5 | 69.4 |  | 70.2 |  | 73.7 |  |  | 21.5 | 24.2 | 23.7 | 25.1 | 26.5 | 69.2 |  | 0 |  | 75.7 |  | 67
[6 empty] | 40 |  | 23.6 |  |  | 26.2 | 24 | 25.2 | 22.8 | 69.5 |  | 70 |  | 73.5 |  |  | 21.8 | 24.8 | 23.2 | 25.3 | 26.5 | 69.1 |  | 0 |  | 75.6 |  | 67
[6 empty] | 40 |  | 24 |  |  | 25.9 | 23.7 | 24.8 | 22.4 | 69.4 |  | 70 |  | 73.5 |  |  | 21.9 | 23.1 | 23.3 | 24.2 | 25.4 | 69.2 |  | 0 |  | 75.8 |  | 67.1
[5 empty] | 3 | 40 | 40 | 24.9 | 24.7 |  | 27.5 | 25.4 | 26 | 24.1 | 69.4 | PASS | 70.1 | PASS | 73.6 | PASS |  | IS1081-IS6110 Ct Mean per Data Set | rpo1 Ct Mean per Data Set | rpo2 Ct Mean per Data Set | rpo3 Ct Mean per Data Set | rpo4 Ct Mean per Data Set
[6 empty] | 40 |  | 24.8 |  |  | 26 | 23.4 | 24.4 | 22.2 | 69.4 |  | 70.1 |  | 73.6 |  |  | 22.8 | 24.8 | 24.7 | 25.9 | 27
[5 empty] | 3 | 40 |  | 25.5 |  |  | 24.7 | 22.5 | 23.2 | 20.7 | 69.2 |  | 69.9 |  | 73.4 |  |  | PASS | PASS | PASS | PASS | PASS
[6 empty] | 40 |  | 25.3 |  |  | 26.6 | 24.1 | 24.8 | 22.5 | 69.3 |  | 70 |  | 73.5
[6 empty] | 40 |  | 24.5 |  | 4 | 25 | 22.5 | 23.4 | 21 | 69.4 | 69.4 | 70.1 | 70 | 73.6 | 73.5
[7 empty] | PASS |  | PASS |  | 25.8 | 23.6 | 24.3 | 22.1 | 69.4 |  | 70 |  | 73.5
[11 empty] | 25.5 | 23.2 | 24.1 | 21.7 | 69.4 |  | 70 |  | 73.5
[6 empty] | 40 |  | 25.1 |  |  | 26.9 | 24.5 | 25.2 | 22.9 | 69.4 |  | 70 |  | 73.5
[6 empty] | 40 |  | 25.1 |  |  | 26.9 | 25.1 | 25.7 | 23.6 | 69.4 | PASS | 69.9 | PASS | 73.4 | PASS
[6 empty] | 40 |  | 22.7 |  |  | 26.4 | 24.7 | 25.7 | 23.2 | 69.5 |  | 70.1 |  | 73.6
[5 empty] | 4 | 40 | 40 | 25.1 | 24.8 |  | 24.8 | 23.1 | 23.7 | 21.4 | 69.3 |  | 69.9 |  | 73.4
[6 empty] | 40 |  | 25.3 |  |  | 26.1 | 23.9 | 24.5 | 22.2 | 69.2 |  | 69.9 |  | 73.4
[6 empty] | 40 |  | 24.4 |  | 5 | 26.1 | 24.1 | 24.2 | 22.3 | 69.4 | 69.4 | 70 | 70 | 73.6 | 73.5
[6 empty] | 40 |  | 23.7 |  |  | 26 | 23.6 | 24.8 | 22.5 | 69.3 |  | 70.1 |  | 73.5
[5 empty] | 4 | 40 |  | 24.4 |  |  | 26.1 | 23.5 | 24.5 | 22.4 | 69.3 |  | 70 |  | 73.5
[6 empty] | 40 | PASS | 23.7 | PASS |  | 25.4 | 23.1 | 24 | 21.7 | 69.5 |  | 69.9 |  | 73.5
[6 empty] | 40 |  | 26.6 |  |  | 25 | 22.5 | 23.9 | 21.3 | 69.3 | PASS | 70 | PASS | 73.6 | PASS
[6 empty] | 40 |  | 24.6 |  |  | 26.6 | 23.6 | 24.6 | 22.3 | 69.3 |  | 69.9 |  | 73.4
[6 empty] | 40 |  | 25.4 |  |  | 25.9 | 23.2 | 24.2 | 22 | 69.4 |  | 70 |  | 73.6
[6 empty] | 40 |  | 24.6 |  |  | 26.3 | 23.5 | 25 | 22.4 | 69.4 |  | 70 |  | 73.6
[5 empty] | 5 | 40 | 40 | 25.7 | 24.9 | 6 | 25.6 | 23.4 | 24.4 | 21.7 | 69.2 | 69.3 | 70 | 70 | 73.5 | 73.5
[6 empty] | 40 |  | 25.8 |  |  | 25.9 | 23.1 | 24.2 | 21.9 | 69.3 |  | 70 |  | 73.5
[6 empty] | 40 |  | 25 |  |  | 25.7 | 22.3 | 23.5 | 21.2 | 69.2 |  | 69.9 |  | 73.5
[6 empty] | 40 |  | 24.4 |  |  | 26.5 | 24 | 25.1 | 22.5 | 69.4 |  | 70.1 |  | 73.5
[6 empty] | 40 |  | 25.9 |  |  | 26.1 | 23.8 | 24.8 | 22.4 | 69.4 | PASS | 70 | PASS | 73.5 | PASS
[5 empty] | 5 | 40 | PASS | 24.1 | PASS |  | 25.5 | 23.1 | 24.2 | 21.9 | 69.3 |  | 70 |  | 73.5
[6 empty] | 40 |  | 23.9 |  |  | 26.1 | 23.9 | 24.9 | 22.7 | 69.4 |  | 70 |  | 73.6
[6 empty] | 40 |  | 24.7 |  |  | 25.8 | 23.3 | 24.3 | 21.9 | 69.5 |  | 70.1 |  | 73.6
[6 empty] | 40 |  | 24.8 |  | 7 | 25.3 | 22.6 | 23.5 | 21.2 | 69.4 | 69.3 | 70 | 70 | 73.6 | 73.5
[6 empty] | 40 |  | 24.8 |  |  | 27.1 | 24.2 | 25.2 | 22.7 | 69.4 |  | 70.1 |  | 73.6
[5 empty] | 6 | 40 | 40 | 25 | 24.5
[6 empty] | 40 |  | 25.8 |  |  | 27.1 | 24.4 | 25.9 | 23.4 | 69.1 |  | 69.8 |  | 73.4
[5 empty] | 6 | 40 |  | 24 |  |  | 26.6 | 24.9 | 25.7 | 23 | 69.3 | PASS | 69.9 | PASS | 73.4 | PASS
[6 empty] | 40 |  | 24.5 |  |  | 26.9 | 24.8 | 25.9 | 23.6 | 69.4 |  | 69.9 |  | 73.5
[6 empty] | 40 |  | 23.7 |  |  | 26 | 23.3 | 24.4 | 22.2 | 69.4 |  | 70.1 |  | 73.6
[6 empty] | 40 | PASS | 25.9 | PASS |  | 26.5 | 23.5 | 24.6 | 22.4 | 69.3 |  | 69.9 |  | 73.4
[6 empty] | 40 |  | 24.2 |  | 8 | 27.7 | 26 | 26.9 | 24.6 | 69.4 | 69.3 | 70.2 | 70 | 73.6 | 73.5
[6 empty] | 40 |  | 24.1 |  |  | 26.6 | 24.6 | 25.8 | 23 | 69.3 |  | 69.9 |  | 73.4
[6 empty] | 40 |  | 24 |  |  | 28 | 25.7 | 26.7 | 24.6 | 69.5 |  | 70 |  | 73.6
[6 empty] | 40 |  | 23.7 |  |  | 26.4 | 23.5 | 24.7 | 22.4 | 69.2 |  | 69.9 |  | 73.4
[5 empty] | 7 | 40 | 40 | 24.5 | 24.6 |  | 27.7 | 26 | 27.4 | 24.8 | 69.4 | PASS | 70.1 | PASS | 73.5 | PASS
[6 empty] | 40 |  | 24.1 |  |  | 25.9 | 24 | 24.8 | 22.4 | 69.2 |  | 69.8 |  | 73.3
[6 empty] | 40 |  | 24.7 |  |  | 26.9 | 24.6 | 25.6 | 23.4 | 69.3 |  | 69.9 |  | 73.5
[6 empty] | 40 |  | 25.7 |  |  | 25.8 | 24.2 | 25.1 | 23 | 69.3 |  | 70 |  | 73.5
[5 empty] | 7 | 40 |  | 24.2 |  | 9 | 25.4 | 23.4 | 24.4 | 21.9 | 69.4 | 69.3 | 70.1 | 69.9 | 73.7 | 73.5
[6 empty] | 40 | PASS | 23.8 | PASS |  | 26.7 | 24.2 | 25.1 | 22.3 | 69.2 |  | 69.9 |  | 73.4
[6 empty] | 40 |  | 24.8 |  |  | 26 | 22.9 | 23.9 | 21.8 | 69.2 |  | 69.8 |  | 73.4
[6 empty] | 40 |  | 26.2 |  |  | 26.5 | 22.9 | 24.3 | 21.7 | 69.1 |  | 69.7 |  | 73.3
[6 empty] | 40 |  | 23.9 |  |  | 25.3 | 23.1 | 23.5 | 20.6 | 69.2 | PASS | 69.8 | PASS | 73.4 | PASS
[6 empty] | 40 |  | 24.1 |  |  | 26.9 | 24.7 | 26 | 23.7 | 69.4 |  | 70.1 |  | 73.6
[5 empty] | 8 | 40 | 40 | 24.8 | 24.3 |  | 25.7 | 23.5 | 24.6 | 22.4 | 69.3 |  | 70 |  | 73.6
[6 empty] | 40 |  | 23.6 |  |  | 26 | 23.3 | 24.5 | 22.2 | 69.3 |  | 70 |  | 73.5
[6 empty] | 40 |  | 25.5 |  | 10 | 25.6 | 23.1 | 24.1 | 21.7 | 69.3 | 69.3 | 70 | 70 | 73.5 | 73.5
[6 empty] | 40 |  | 26.1 |  |  | 26.8 | 24.6 | 25.4 | 23.1 | 69.4 |  | 70 |  | 73.6
[6 empty] | 40 |  | 23.4 |  |  | 25.4 | 23.2 | 24.4 | 22 | 69.4 |  | 70.1 |  | 73.6
[6 empty] | 40 | PASS | 24.7 | PASS |  | 26 | 23.2 | 24.4 | 22.2 | 69.4 |  | 70.1 |  | 73.6
[5 empty] | 8 | 40 |  | 24.1 | [7 empty] | PASS |  | PASS |  | PASS
[6 empty] | 40 |  | 23.2 |  |  | 25.9 | 23.6 | 24.4 | 22.2 | 69.3 |  | 70 |  | 73.5
[6 empty] | 40 |  | 23.5 |  |  | 25.4 | 23.4 | 24 | 21.6 | 69.2 |  | 69.8 |  | 73.3
[6 empty] | 40 |  | 23.8 |  |  | 25.5 | 23.2 | 23.9 | 21.5 | 69.3 |  | 69.9 |  | 73.5
[5 empty] | 9 | 40 | 40 | 24.2 | 24.5 |  | IS1081-IS6110 Ct Mean per Data Set | rpo1 Ct Mean per Data Set | rpo2 Ct Mean per Data Set | rpo4 Ct Mean per Data Set
[6 empty] | 40 |  | 23.9 |  |  | 26.1 | 23.8 | 24.7 | 22.4
[6 empty] | 40 |  | 25 |  |  | PASS | PASS | PASS | PASS
[5 empty] | 9 | 40 |  | 26
[6 empty] | 40 |  | 25
[6 empty] | 40 | PASS | 25 | PASS
[6 empty] | 40 |  | 23.8
[6 empty] | 40 |  | 24.1
[6 empty] | 40 |  | 24.4
[6 empty] | 40 |  | 23.9
[5 empty] | 10 | 40 | 40 | 23.8 | 24.3
[6 empty] | 40 |  | 24.6
[6 empty] | 40 |  | 25.1
[6 empty] | 40 |  | 24.4
[5 empty] | 10 | 40 |  | 23.9
[7 empty] | PASS |  | PASS
[6 empty] | 40 |  | 23.7
[6 empty] | 40 |  | 26.4
[6 empty] | 40 |  | 22.8
[6 empty] | 40 |  | 23.9
FIO

## Data - First 10
Assay | Assay Version | Sample ID | Cartridge S/N | Reagent Lot ID | Test Type | Instrument S/N | Module S/N | S/W Version | Start Time | Status | Test Result | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Max Pressure (PSI) | Min Pressure (PSI) | Error | Patient ID | Quantitative Value | Quantitative Unit | Interpretation Flag | User
Xpert MTB-RIF Ultra | 4 | NEG_00_0_01X | 614332136 | 92303 | Negative Control 1 | 803094 | 730647 | 6.8 | 2025-10-26T01:27:08+0200 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 24.3 | 147 | PASS | -9999 | 226 | 496 | 219 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 5 | NA | -9999 | 223 | 527 | 219 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 1 | NA | -9999 | 281 | 693 | 277 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 0 | NA | -9999 | 114 | 346 | 112 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 7 | NA | -9999 | 191 | 457 | 191 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 4 | NA | -9999 | 153 | 403 | 150 | PASS | 0 | 40 | QC1 | Probe | 125 | 124 | 124 | PASS | QC2 | Probe | 129 | 129 | 128 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 18.9 | -7.5 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_00_0_02X | 614332146 | 92303 | Negative Control 1 | 803094 | 773527 | 6.8 | 2025-10-26T01:27:35+0200 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 24.3 | 151 | PASS | -9999 | 213 | 475 | 208 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 8 | NA | -9999 | 213 | 506 | 211 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 9 | NA | -9999 | 271 | 679 | 270 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 4 | NA | -9999 | 104 | 320 | 103 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 6 | NA | -9999 | 197 | 469 | 196 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 4 | NA | -9999 | 150 | 394 | 147 | PASS | 0 | 40 | QC1 | Probe | 120 | 119 | 119 | PASS | QC2 | Probe | 122 | 122 | 122 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 19 | -6.7 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_00_0_03X | 614332148 | 92303 | Negative Control 1 | 803094 | 210043140 | 6.8 | 2025-10-26T01:28:05+0200 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 24.5 | 151 | PASS | -9999 | 215 | 466 | 208 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 6 | NA | -9999 | 204 | 486 | 201 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 0 | NA | -9999 | 284 | 695 | 281 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 0 | NA | -9999 | 105 | 331 | 102 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 7 | NA | -9999 | 180 | 429 | 179 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 9 | NA | -9999 | 141 | 366 | 139 | PASS | 0 | 40 | QC1 | Probe | 119 | 119 | 118 | PASS | QC2 | Probe | 118 | 117 | 117 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 19.4 | -7.7 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_00_0_04X | 614332134 | 92303 | Negative Control 1 | 803094 | 909441 | 6.8 | 2025-10-26T01:28:32+0200 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 23.2 | 158 | PASS | -9999 | 181 | 377 | 172 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 8 | NA | -9999 | 183 | 438 | 180 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 2 | NA | -9999 | 246 | 611 | 244 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 7 | NA | -9999 | 95 | 302 | 93 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 11 | NA | -9999 | 162 | 367 | 161 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 9 | NA | -9999 | 116 | 304 | 114 | PASS | 0 | 40 | QC1 | Probe | 103 | 102 | 102 | PASS | QC2 | Probe | 114 | 114 | 114 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 17.5 | -5.8 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_00_0_05X | 614332132 | 92303 | Negative Control 1 | 803094 | 880243 | 6.8 | 2025-10-26T01:29:09+0200 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 21.6 | 177 | PASS | -9999 | 212 | 483 | 207 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 9 | NA | -9999 | 205 | 490 | 202 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 25 | NA | -9999 | 287 | 717 | 283 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 5 | NA | -9999 | 100 | 322 | 99 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 10 | NA | -9999 | 179 | 415 | 177 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 10 | NA | -9999 | 118 | 312 | 117 | PASS | 0 | 40 | QC1 | Probe | 126 | 125 | 125 | PASS | QC2 | Probe | 129 | 129 | 129 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 14.8 | -9.4 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_00_0_06+ | 614332137 | 92303 | Negative Control 1 | 803094 | 210013102 | 6.8 | 2025-10-26T01:29:41+0200 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 23.9 | 148 | PASS | -9999 | 187 | 411 | 181 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 9 | NA | -9999 | 190 | 456 | 186 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 6 | NA | -9999 | 240 | 598 | 238 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 3 | NA | -9999 | 97 | 307 | 95 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 8 | NA | -9999 | 168 | 391 | 166 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 5 | NA | -9999 | 123 | 321 | 120 | PASS | 0 | 40 | QC1 | Probe | 98 | 97 | 98 | PASS | QC2 | Probe | 109 | 108 | 108 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 16.9 | -9.3 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_00_0_07+ | 614332149 | 92303 | Negative Control 1 | 803094 | 674169 | 6.8 | 2025-10-26T01:30:04+0200 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 24.5 | 145 | PASS | -9999 | 212 | 464 | 205 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 4 | NA | -9999 | 216 | 514 | 213 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 2 | NA | -9999 | 282 | 693 | 278 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 1 | NA | -9999 | 119 | 358 | 116 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 8 | NA | -9999 | 174 | 402 | 172 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 4 | NA | -9999 | 146 | 381 | 143 | PASS | 0 | 40 | QC1 | Probe | 125 | 124 | 124 | PASS | QC2 | Probe | 119 | 119 | 119 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 17.3 | -8.8 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_00_0_08+ | 614332145 | 92303 | Negative Control 1 | 803094 | 631549 | 6.8 | 2025-10-26T01:30:36+0200 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 24.6 | 152 | PASS | -9999 | 224 | 489 | 216 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 11 | NA | -9999 | 196 | 465 | 194 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 5 | NA | -9999 | 275 | 679 | 272 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | -1 | NA | -9999 | 109 | 343 | 108 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 9 | NA | -9999 | 167 | 390 | 167 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 6 | NA | -9999 | 141 | 365 | 138 | PASS | 0 | 40 | QC1 | Probe | 120 | 120 | 119 | PASS | QC2 | Probe | 100 | 100 | 100 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 22.5 | -12.3 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_00_0_09+ | 614332151 | 92303 | Negative Control 1 | 803094 | 736909 | 6.8 | 2025-10-26T01:31:03+0200 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 24.4 | 146 | PASS | -9999 | 199 | 435 | 190 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 6 | NA | -9999 | 197 | 470 | 195 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 3 | NA | -9999 | 269 | 668 | 266 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 3 | NA | -9999 | 97 | 306 | 95 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 5 | NA | -9999 | 161 | 392 | 162 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 10 | NA | -9999 | 148 | 379 | 145 | PASS | 0 | 40 | QC1 | Probe | 124 | 124 | 124 | PASS | QC2 | Probe | 110 | 110 | 109 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 18.3 | -6.9 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_00_0_10+ | 614332135 | 92303 | Negative Control 1 | 803094 | 914633 | 6.8 | 2025-10-26T01:31:27+0200 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 25 | 152 | PASS | -9999 | 218 | 471 | 212 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 9 | NA | -9999 | 223 | 530 | 220 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 13 | NA | -9999 | 285 | 708 | 282 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 7 | NA | -9999 | 110 | 342 | 108 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 7 | NA | -9999 | 183 | 434 | 182 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 8 | NA | -9999 | 149 | 386 | 146 | PASS | 0 | 40 | QC1 | Probe | 120 | 120 | 119 | PASS | QC2 | Probe | 116 | 116 | 116 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 18.6 | -8.6 | [5 empty] | <None>

## Data - Sample Type 0
Assay | Assay Version | Sample ID | Cartridge S/N | Reagent Lot ID | Test Type | Instrument S/N | Module S/N | S/W Version | Start Time | Status | Test Result | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Max Pressure (PSI) | Min Pressure (PSI) | Error | Patient ID | Quantitative Value | Quantitative Unit | Interpretation Flag | User
Xpert MTB-RIF Ultra | 4 | NEG_01_0_01X | 614332262 | 92303 | Negative Control 1 | 847922 | 604077 | 6.8 | 2025-10-26T02:56:07+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 26.3 | 128 | PASS | -9999 | 193 | 420 | 190 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 6 | NA | -9999 | 198 | 473 | 197 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 11 | NA | -9999 | 235 | 572 | 234 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | -1 | NA | -9999 | 104 | 321 | 103 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | -1 | NA | -9999 | 173 | 411 | 173 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 4 | NA | -9999 | 154 | 403 | 152 | PASS | 0 | 40 | QC1 | Probe | 119 | 118 | 118 | PASS | QC2 | Probe | 112 | 112 | 111 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 19.2 | -10.1 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_01_0_02X | 614332258 | 92303 | Negative Control 1 | 847922 | 210015588 | 6.8 | 2025-10-26T02:56:32+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 24 | 143 | PASS | -9999 | 209 | 461 | 205 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 6 | NA | -9999 | 171 | 409 | 169 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 0 | NA | -9999 | 237 | 592 | 235 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 3 | NA | -9999 | 102 | 312 | 99 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 3 | NA | -9999 | 188 | 452 | 188 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 4 | NA | -9999 | 129 | 337 | 127 | PASS | 0 | 40 | QC1 | Probe | 102 | 101 | 101 | PASS | QC2 | Probe | 115 | 114 | 114 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 21.6 | -10.2 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_01_0_03X | 614332256 | 92303 | Negative Control 1 | 847922 | 660764 | 6.8 | 2025-10-26T02:56:56+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 23.9 | 135 | PASS | -9999 | 186 | 404 | 181 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 5 | NA | -9999 | 199 | 470 | 195 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 22 | NA | -9999 | 266 | 646 | 261 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 0 | NA | -9999 | 110 | 333 | 108 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 5 | NA | -9999 | 188 | 423 | 185 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 6 | NA | -9999 | 152 | 396 | 149 | PASS | 0 | 40 | QC1 | Probe | 125 | 125 | 125 | PASS | QC2 | Probe | 129 | 129 | 128 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 15.6 | -5.7 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_01_0_04X | 614332260 | 92303 | Negative Control 1 | 847922 | 617583 | 6.8 | 2025-10-26T02:57:42+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 23.7 | 133 | PASS | -9999 | 209 | 444 | 197 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 3 | NA | -9999 | 197 | 469 | 191 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | -1 | NA | -9999 | 265 | 651 | 260 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 0 | NA | -9999 | 90 | 279 | 87 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 2 | NA | -9999 | 172 | 397 | 170 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 2 | NA | -9999 | 146 | 389 | 142 | PASS | 0 | 40 | QC1 | Probe | 126 | 126 | 125 | PASS | QC2 | Probe | 112 | 112 | 111 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 18.5 | -9.8 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_01_0_05X | 614332264 | 92303 | Negative Control 1 | 847922 | 725382 | 6.8 | 2025-10-26T02:58:58+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 27.1 | 123 | PASS | -9999 | 196 | 426 | 193 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 4 | NA | -9999 | 204 | 484 | 202 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 6 | NA | -9999 | 272 | 668 | 270 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 1 | NA | -9999 | 99 | 302 | 97 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 9 | NA | -9999 | 164 | 367 | 163 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 1 | NA | -9999 | 143 | 382 | 140 | PASS | 0 | 40 | QC1 | Probe | 123 | 123 | 122 | PASS | QC2 | Probe | 109 | 109 | 109 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 21.5 | -9.5 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_01_0_06+ | 614332263 | 92303 | Negative Control 1 | 847922 | 710982 | 6.8 | 2025-10-26T02:59:48+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 23.6 | 139 | PASS | -9999 | 217 | 475 | 212 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 8 | NA | -9999 | 188 | 449 | 186 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | -6 | NA | -9999 | 262 | 650 | 261 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 3 | NA | -9999 | 101 | 320 | 99 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 3 | NA | -9999 | 199 | 475 | 200 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 3 | NA | -9999 | 139 | 362 | 137 | PASS | 0 | 40 | QC1 | Probe | 121 | 121 | 121 | PASS | QC2 | Probe | 128 | 127 | 127 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 22 | -12.1 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_01_0_07+ | 614332261 | 92303 | Negative Control 1 | 847922 | 912152 | 6.8 | 2025-10-26T02:00:11+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 23.3 | 149 | PASS | -9999 | 204 | 452 | 195 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 5 | NA | -9999 | 209 | 493 | 206 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 1 | NA | -9999 | 269 | 664 | 265 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 1 | NA | -9999 | 103 | 335 | 101 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 2 | NA | -9999 | 182 | 432 | 181 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 4 | NA | -9999 | 144 | 384 | 141 | PASS | 0 | 40 | QC1 | Probe | 123 | 123 | 122 | PASS | QC2 | Probe | 120 | 120 | 120 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 16 | -12.7 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_01_0_08A+ | 614332245 | 92303 | Negative Control 1 | 847922 | 730163 | 6.8 | 2025-10-26T02:34:57+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 25.8 | 137 | PASS | -9999 | 191 | 404 | 187 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 3 | NA | -9999 | 175 | 413 | 174 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | -11 | NA | -9999 | 248 | 622 | 246 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 0 | NA | -9999 | 93 | 286 | 91 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 2 | NA | -9999 | 147 | 352 | 148 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 6 | NA | -9999 | 109 | 298 | 108 | PASS | 0 | 40 | QC1 | Probe | 90 | 90 | 90 | PASS | QC2 | Probe | 91 | 91 | 91 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 15.5 | -12.3 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_01_0_09+ | 614332267 | 92303 | Negative Control 1 | 847922 | 892236 | 6.8 | 2025-10-26T02:00:36+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 25 | 147 | PASS | -9999 | 213 | 466 | 208 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 5 | NA | -9999 | 216 | 515 | 214 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 3 | NA | -9999 | 268 | 662 | 266 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 4 | NA | -9999 | 110 | 331 | 109 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 4 | NA | -9999 | 195 | 468 | 193 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 6 | NA | -9999 | 133 | 344 | 130 | PASS | 0 | 40 | QC1 | Probe | 106 | 105 | 105 | PASS | QC2 | Probe | 116 | 116 | 115 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 18.9 | -7.3 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_01_0_10+ | 614332259 | 92303 | Negative Control 1 | 847922 | 210096147 | 6.8 | 2025-10-26T02:01:01+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 24.2 | 143 | PASS | -9999 | 184 | 394 | 179 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 5 | NA | -9999 | 157 | 372 | 155 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 8 | NA | -9999 | 212 | 520 | 210 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 3 | NA | -9999 | 99 | 298 | 97 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 6 | NA | -9999 | 164 | 363 | 163 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 8 | NA | -9999 | 112 | 289 | 110 | PASS | 0 | 40 | QC1 | Probe | 95 | 94 | 94 | PASS | QC2 | Probe | 113 | 113 | 112 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 17.1 | -8 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_02_0_01X | 614333120 | 92303 | Negative Control 1 | 847922 | 897302 | 6.8 | 2025-10-26T03:39:14+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 24.4 | 140 | PASS | -9999 | 260 | 561 | 255 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 6 | NA | -9999 | 199 | 476 | 197 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 3 | NA | -9999 | 274 | 674 | 272 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 1 | NA | -9999 | 127 | 392 | 124 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 6 | NA | -9999 | 174 | 422 | 174 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 7 | NA | -9999 | 155 | 403 | 153 | PASS | 0 | 40 | QC1 | Probe | 104 | 104 | 104 | PASS | QC2 | Probe | 112 | 112 | 112 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 17.2 | -8.1 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_02_0_02X | 614332906 | 92303 | Negative Control 1 | 847922 | 210097747 | 6.8 | 2025-10-26T03:39:40+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 24.4 | 146 | PASS | -9999 | 191 | 420 | 185 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 2 | NA | -9999 | 193 | 458 | 191 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 1 | NA | -9999 | 253 | 626 | 252 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 0 | NA | -9999 | 107 | 331 | 105 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 8 | NA | -9999 | 166 | 379 | 166 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 7 | NA | -9999 | 132 | 343 | 130 | PASS | 0 | 40 | QC1 | Probe | 121 | 121 | 120 | PASS | QC2 | Probe | 121 | 121 | 121 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 17 | -4.5 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_02_0_03X | 614333486 | 92303 | Negative Control 1 | 847922 | 210013071 | 6.8 | 2025-10-26T03:40:05+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 24.4 | 156 | PASS | -9999 | 210 | 460 | 205 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 5 | NA | -9999 | 201 | 480 | 199 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 6 | NA | -9999 | 263 | 653 | 260 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 3 | NA | -9999 | 110 | 340 | 108 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 10 | NA | -9999 | 158 | 387 | 158 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 6 | NA | -9999 | 144 | 377 | 142 | PASS | 0 | 40 | QC1 | Probe | 113 | 112 | 112 | PASS | QC2 | Probe | 99 | 99 | 99 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 19.6 | -8.3 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_02_0_04X | 614333754 | 92303 | Negative Control 1 | 847922 | 657194 | 6.8 | 2025-10-26T03:40:54+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 24.1 | 162 | PASS | -9999 | 217 | 474 | 207 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 4 | NA | -9999 | 190 | 451 | 187 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 1 | NA | -9999 | 260 | 641 | 256 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 5 | NA | -9999 | 109 | 328 | 107 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 9 | NA | -9999 | 150 | 363 | 150 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 3 | NA | -9999 | 144 | 375 | 142 | PASS | 0 | 40 | QC1 | Probe | 114 | 113 | 113 | PASS | QC2 | Probe | 92 | 92 | 92 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 21.1 | -9 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_02_0_05X | 614332568 | 92303 | Negative Control 1 | 847922 | 702829 | 6.8 | 2025-10-26T03:41:19+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 24.2 | 154 | PASS | -9999 | 235 | 523 | 229 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 7 | NA | -9999 | 207 | 492 | 205 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 10 | NA | -9999 | 261 | 645 | 259 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 2 | NA | -9999 | 120 | 366 | 119 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 6 | NA | -9999 | 182 | 438 | 182 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 5 | NA | -9999 | 158 | 414 | 156 | PASS | 0 | 40 | QC1 | Probe | 139 | 138 | 138 | PASS | QC2 | Probe | 119 | 119 | 119 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 17.6 | -7.9 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_02_0_06+ | 614333339 | 92303 | Negative Control 1 | 847922 | 750444 | 6.8 | 2025-10-26T03:41:42+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 28.3 | 128 | PASS | -9999 | 217 | 472 | 208 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 2 | NA | -9999 | 199 | 477 | 197 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | -9 | NA | -9999 | 255 | 624 | 253 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 1 | NA | -9999 | 111 | 344 | 110 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 6 | NA | -9999 | 172 | 404 | 170 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 8 | NA | -9999 | 145 | 376 | 143 | PASS | 0 | 40 | QC1 | Probe | 111 | 111 | 112 | PASS | QC2 | Probe | 96 | 96 | 95 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 18.5 | -7.4 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_02_0_07+ | 614332989 | 92303 | Negative Control 1 | 847922 | 220000161 | 6.8 | 2025-10-26T03:42:06+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 25.3 | 133 | PASS | -9999 | 229 | 492 | 218 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 8 | NA | -9999 | 188 | 448 | 186 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | -7 | NA | -9999 | 270 | 668 | 266 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 0 | NA | -9999 | 107 | 319 | 105 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 8 | NA | -9999 | 155 | 365 | 154 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 2 | NA | -9999 | 133 | 348 | 130 | PASS | 0 | 40 | QC1 | Probe | 107 | 106 | 106 | PASS | QC2 | Probe | 97 | 96 | 96 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 19.8 | -8.2 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_02_0_08+ | 614332693 | 92303 | Negative Control 1 | 847922 | 760491 | 6.8 | 2025-10-26T03:42:31+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 24.5 | 140 | PASS | -9999 | 216 | 470 | 209 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 7 | NA | -9999 | 195 | 465 | 192 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 4 | NA | -9999 | 252 | 623 | 249 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 3 | NA | -9999 | 109 | 333 | 106 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 4 | NA | -9999 | 184 | 449 | 184 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 5 | NA | -9999 | 143 | 372 | 140 | PASS | 0 | 40 | QC1 | Probe | 114 | 113 | 113 | PASS | QC2 | Probe | 124 | 124 | 124 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 19.9 | -8.1 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_02_0_09+ | 614333581 | 92303 | Negative Control 1 | 847922 | 210068154 | 6.8 | 2025-10-26T03:42:56+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 23.6 | 145 | PASS | -9999 | 208 | 453 | 202 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 4 | NA | -9999 | 184 | 438 | 183 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 0 | NA | -9999 | 255 | 637 | 253 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 2 | NA | -9999 | 105 | 327 | 104 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 5 | NA | -9999 | 186 | 445 | 187 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 7 | NA | -9999 | 127 | 329 | 125 | PASS | 0 | 40 | QC1 | Probe | 112 | 111 | 111 | PASS | QC2 | Probe | 128 | 128 | 128 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 18.2 | -8.1 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_02_0_10+ | 614332391 | 92303 | Negative Control 1 | 847922 | 902140 | 6.8 | 2025-10-26T03:43:21+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 24 | 156 | PASS | -9999 | 183 | 401 | 177 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 6 | NA | -9999 | 183 | 437 | 182 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 3 | NA | -9999 | 243 | 604 | 241 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 4 | NA | -9999 | 89 | 275 | 88 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 9 | NA | -9999 | 151 | 366 | 151 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 4 | NA | -9999 | 140 | 370 | 138 | PASS | 0 | 40 | QC1 | Probe | 113 | 113 | 112 | PASS | QC2 | Probe | 100 | 100 | 100 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 15.7 | -9.4 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_03_0_01X | 614337122 | 92303 | Negative Control 1 | 847922 | 897302 | 6.8 | 2025-10-26T05:52:14+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 24.9 | 149 | PASS | -9999 | 247 | 534 | 240 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 6 | NA | -9999 | 196 | 469 | 194 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 5 | NA | -9999 | 264 | 654 | 262 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 0 | NA | -9999 | 127 | 386 | 125 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 6 | NA | -9999 | 166 | 402 | 166 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 7 | NA | -9999 | 130 | 338 | 128 | PASS | 0 | 40 | QC1 | Probe | 103 | 103 | 103 | PASS | QC2 | Probe | 104 | 104 | 104 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 12.7 | -11.3 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_03_0_02X | 614335906 | 92303 | Negative Control 1 | 847922 | 210097747 | 6.8 | 2025-10-26T05:52:40+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 24.8 | 155 | PASS | -9999 | 195 | 431 | 188 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 3 | NA | -9999 | 193 | 455 | 190 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 4 | NA | -9999 | 256 | 628 | 253 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | -1 | NA | -9999 | 108 | 334 | 106 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 8 | NA | -9999 | 165 | 379 | 164 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 7 | NA | -9999 | 135 | 346 | 132 | PASS | 0 | 40 | QC1 | Probe | 113 | 113 | 112 | PASS | QC2 | Probe | 104 | 104 | 104 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 22.3 | -10.7 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_03_0_03X | 614336252 | 92303 | Negative Control 1 | 847922 | 210013071 | 6.8 | 2025-10-26T05:53:05+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 25.5 | 144 | PASS | -9999 | 190 | 399 | 185 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 7 | NA | -9999 | 176 | 417 | 174 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 2 | NA | -9999 | 229 | 558 | 227 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 1 | NA | -9999 | 101 | 308 | 99 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 16 | NA | -9999 | 143 | 313 | 142 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 5 | NA | -9999 | 142 | 366 | 140 | PASS | 0 | 40 | QC1 | Probe | 111 | 111 | 111 | PASS | QC2 | Probe | 97 | 96 | 96 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 18.8 | -7.1 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_03_0_04X | 614336830 | 92303 | Negative Control 1 | 847922 | 657194 | 6.8 | 2025-10-26T05:53:30+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 25.3 | 149 | PASS | -9999 | 220 | 479 | 211 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 3 | NA | -9999 | 193 | 457 | 191 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | -1 | NA | -9999 | 267 | 664 | 265 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 2 | NA | -9999 | 108 | 328 | 105 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 7 | NA | -9999 | 156 | 375 | 157 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 1 | NA | -9999 | 146 | 382 | 144 | PASS | 0 | 40 | QC1 | Probe | 104 | 104 | 104 | PASS | QC2 | Probe | 95 | 95 | 95 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 18.7 | -10.8 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_03_0_05X | 614336556 | 92303 | Negative Control 1 | 847922 | 702829 | 6.8 | 2025-10-26T05:53:53+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 24.5 | 142 | PASS | -9999 | 229 | 503 | 225 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 6 | NA | -9999 | 189 | 451 | 188 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 7 | NA | -9999 | 241 | 597 | 240 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 2 | NA | -9999 | 114 | 345 | 112 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 3 | NA | -9999 | 177 | 416 | 176 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 3 | NA | -9999 | 158 | 412 | 156 | PASS | 0 | 40 | QC1 | Probe | 135 | 135 | 135 | PASS | QC2 | Probe | 114 | 114 | 113 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 17.3 | -7.9 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_03_0_06+ | 614336421 | 92303 | Negative Control 1 | 847922 | 750444 | 6.8 | 2025-10-26T05:54:16+0100 | Done | MTB Trace DETECTED;RIF Resistance INDETERMINATE | SPC | SPC | 1,2,3,4 | PASS | 24.4 | 142 | PASS | -9999 | 226 | 491 | 223 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | PASS | 27.4 | 461 | PASS | -9999 | 220 | 529 | 221 | PASS | 0 | 40 | rpoB1 | Target | 1 | NEG | 0 | 2 | NA | -9999 | 263 | 655 | 265 | PASS | 0 | 40 | rpoB2 | Target | 2 | NEG | 0 | 3 | NA | -9999 | 116 | 365 | 116 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 6 | NA | -9999 | 178 | 422 | 180 | PASS | 0 | 40 | rpoB4 | Target | 4 | NEG | 0 | 6 | NA | -9999 | 154 | 406 | 154 | PASS | 0 | 40 | QC1 | Probe | 121 | 121 | 121 | PASS | QC2 | Probe | 106 | 106 | 107 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 13.1 | -9.2 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_03_0_07+ | 614335759 | 92303 | Negative Control 1 | 847922 | 220000161 | 6.8 | 2025-10-26T05:54:41+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | PASS | 24.7 | 147 | PASS | -9999 | 219 | 472 | 208 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 7 | NA | -9999 | 198 | 471 | 196 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | -1 | NA | -9999 | 266 | 658 | 262 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 3 | NA | -9999 | 110 | 333 | 108 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 8 | NA | -9999 | 151 | 361 | 150 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 5 | NA | -9999 | 144 | 380 | 141 | PASS | 0 | 40 | QC1 | Probe | 120 | 120 | 119 | PASS | QC2 | Probe | 103 | 103 | 103 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 16.6 | -6.6 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_03_0_08+ | 614336979 | 92303 | Negative Control 1 | 847922 | 760491 | 6.8 | 2025-10-26T05:55:05+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 25.1 | 128 | PASS | -9999 | 217 | 481 | 211 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 6 | NA | -9999 | 187 | 445 | 185 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | -7 | NA | -9999 | 254 | 634 | 253 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 2 | NA | -9999 | 107 | 326 | 105 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 1 | NA | -9999 | 190 | 466 | 190 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 5 | NA | -9999 | 131 | 342 | 129 | PASS | 0 | 40 | QC1 | Probe | 112 | 112 | 112 | PASS | QC2 | Probe | 123 | 122 | 122 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 17.9 | -7.2 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_03_0_09+ | 614336667 | 92303 | Negative Control 1 | 847922 | 210068154 | 6.8 | 2025-10-26T05:55:30+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 25.1 | 137 | PASS | -9999 | 197 | 428 | 193 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 4 | NA | -9999 | 179 | 427 | 179 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | -5 | NA | -9999 | 248 | 618 | 247 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 3 | NA | -9999 | 104 | 318 | 103 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 3 | NA | -9999 | 178 | 425 | 179 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 7 | NA | -9999 | 126 | 327 | 125 | PASS | 0 | 40 | QC1 | Probe | 112 | 112 | 112 | PASS | QC2 | Probe | 116 | 116 | 116 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 17.7 | -7.9 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_03_0_10+ | 614336051 | 92303 | Negative Control 1 | 847922 | 902140 | 6.8 | 2025-10-26T05:55:53+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 22.7 | 170 | PASS | -9999 | 187 | 404 | 180 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 9 | NA | -9999 | 187 | 448 | 185 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 15 | NA | -9999 | 239 | 598 | 239 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 6 | NA | -9999 | 88 | 274 | 87 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 9 | NA | -9999 | 158 | 365 | 157 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 6 | NA | -9999 | 137 | 364 | 135 | PASS | 0 | 40 | QC1 | Probe | 123 | 122 | 122 | PASS | QC2 | Probe | 114 | 114 | 114 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 16.2 | -6.5 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_04_0_01X | 614340140 | 92303 | Negative Control 1 | 847922 | 210097747 | 6.8 | 2025-10-26T09:34:16+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 25.1 | 137 | PASS | -9999 | 188 | 418 | 182 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 0 | NA | -9999 | 187 | 443 | 185 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | -2 | NA | -9999 | 248 | 608 | 246 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 0 | NA | -9999 | 109 | 334 | 107 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 8 | NA | -9999 | 160 | 370 | 159 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 6 | NA | -9999 | 130 | 336 | 127 | PASS | 0 | 40 | QC1 | Probe | 109 | 108 | 108 | PASS | QC2 | Probe | 102 | 102 | 102 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 19.4 | -6.8 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_04_0_02X | 614339872 | 92303 | Negative Control 1 | 847922 | 210013071 | 6.8 | 2025-10-26T09:34:40+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 25.3 | 149 | PASS | -9999 | 203 | 437 | 198 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 4 | NA | -9999 | 192 | 460 | 190 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 3 | NA | -9999 | 249 | 617 | 247 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 2 | NA | -9999 | 109 | 332 | 106 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 13 | NA | -9999 | 154 | 361 | 153 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 6 | NA | -9999 | 139 | 365 | 137 | PASS | 0 | 40 | QC1 | Probe | 113 | 112 | 112 | PASS | QC2 | Probe | 104 | 103 | 102 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 17 | -9.1 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_04_0_03X | 614339510 | 92303 | Negative Control 1 | 847922 | 657194 | 6.8 | 2025-10-26T09:35:05+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 24.4 | 165 | PASS | -9999 | 221 | 488 | 212 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 5 | NA | -9999 | 196 | 463 | 194 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 3 | NA | -9999 | 267 | 661 | 264 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 2 | NA | -9999 | 111 | 332 | 109 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 10 | NA | -9999 | 152 | 374 | 152 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 4 | NA | -9999 | 142 | 369 | 140 | PASS | 0 | 40 | QC1 | Probe | 107 | 106 | 106 | PASS | QC2 | Probe | 102 | 101 | 101 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 18.2 | -10.6 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_04_0_04X | 614340440 | 92303 | Negative Control 1 | 847922 | 702829 | 6.8 | 2025-10-26T09:35:30+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 23.7 | 159 | PASS | -9999 | 235 | 523 | 231 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 5 | NA | -9999 | 209 | 499 | 208 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 10 | NA | -9999 | 263 | 653 | 262 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 2 | NA | -9999 | 118 | 361 | 117 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 6 | NA | -9999 | 183 | 438 | 182 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 4 | NA | -9999 | 159 | 418 | 158 | PASS | 0 | 40 | QC1 | Probe | 144 | 144 | 143 | PASS | QC2 | Probe | 130 | 129 | 129 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 15.9 | -7.4 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_04_0_05X | 614339216 | 92303 | Negative Control 1 | 847922 | 220000161 | 6.8 | 2025-10-26T09:35:54+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 24.4 | 141 | PASS | -9999 | 229 | 493 | 218 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 7 | NA | -9999 | 204 | 480 | 201 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | -1 | NA | -9999 | 270 | 665 | 266 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | -1 | NA | -9999 | 111 | 328 | 109 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 9 | NA | -9999 | 152 | 369 | 151 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 3 | NA | -9999 | 144 | 376 | 141 | PASS | 0 | 40 | QC1 | Probe | 119 | 118 | 117 | PASS | QC2 | Probe | 101 | 100 | 100 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 16.8 | -8.7 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_04_0_06+ | 614339977 | 92303 | Negative Control 1 | 847922 | 760491 | 6.8 | 2025-10-26T09:36:20+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 23.7 | 150 | PASS | -9999 | 210 | 460 | 203 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 3 | NA | -9999 | 197 | 468 | 193 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 1 | NA | -9999 | 239 | 592 | 236 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 2 | NA | -9999 | 109 | 333 | 106 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 5 | NA | -9999 | 185 | 445 | 185 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 4 | NA | -9999 | 139 | 363 | 137 | PASS | 0 | 40 | QC1 | Probe | 99 | 99 | 99 | PASS | QC2 | Probe | 130 | 130 | 129 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 19.2 | -8.1 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_04_0_07+ | 614340317 | 92303 | Negative Control 1 | 847922 | 210068154 | 6.8 | 2025-10-26T09:36:44+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 26.6 | 122 | PASS | -9999 | 202 | 440 | 198 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 4 | NA | -9999 | 184 | 437 | 182 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | -4 | NA | -9999 | 252 | 626 | 250 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 1 | NA | -9999 | 107 | 327 | 106 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 1 | NA | -9999 | 177 | 431 | 177 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 7 | NA | -9999 | 130 | 336 | 128 | PASS | 0 | 40 | QC1 | Probe | 117 | 117 | 116 | PASS | QC2 | Probe | 90 | 90 | 90 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 17.6 | -8.5 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_04_0_08+ | 614339723 | 92303 | Negative Control 1 | 847922 | 701892 | 6.8 | 2025-10-26T09:37:09+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 24.6 | 144 | PASS | -9999 | 218 | 484 | 212 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 0 | NA | -9999 | 216 | 510 | 213 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | -4 | NA | -9999 | 264 | 653 | 261 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 1 | NA | -9999 | 112 | 337 | 110 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 8 | NA | -9999 | 188 | 456 | 187 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 6 | NA | -9999 | 148 | 388 | 146 | PASS | 0 | 40 | QC1 | Probe | 128 | 127 | 126 | PASS | QC2 | Probe | 120 | 120 | 120 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 17.1 | -7.8 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_04_0_09+ | 614339343 | 92303 | Negative Control 1 | 847922 | 866331 | 6.8 | 2025-10-26T09:37:32+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 25.4 | 153 | PASS | -9999 | 215 | 455 | 204 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 9 | NA | -9999 | 187 | 446 | 184 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 14 | NA | -9999 | 226 | 548 | 223 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 2 | NA | -9999 | 97 | 306 | 94 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 5 | NA | -9999 | 179 | 414 | 177 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 6 | NA | -9999 | 137 | 357 | 134 | PASS | 0 | 40 | QC1 | Probe | 110 | 109 | 109 | PASS | QC2 | Probe | 106 | 105 | 105 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 17.1 | -10.4 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_04_0_10+ | 614338967 | 92303 | Negative Control 1 | 847922 | 902140 | 6.8 | 2025-10-26T09:37:56+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 24.6 | 153 | PASS | -9999 | 194 | 420 | 188 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 7 | NA | -9999 | 192 | 457 | 189 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 7 | NA | -9999 | 248 | 614 | 245 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 4 | NA | -9999 | 92 | 282 | 90 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 7 | NA | -9999 | 156 | 373 | 156 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 3 | NA | -9999 | 146 | 384 | 143 | PASS | 0 | 40 | QC1 | Probe | 117 | 116 | 116 | PASS | QC2 | Probe | 100 | 100 | 100 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 20.5 | -9.9 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_05_0_01X | 614343144 | 92303 | Negative Control 1 | 847922 | 775845 | 6.8 | 2025-10-26T09:42:20+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 25.7 | 146 | PASS | -9999 | 226 | 480 | 221 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 2 | NA | -9999 | 207 | 494 | 206 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | -5 | NA | -9999 | 277 | 682 | 276 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 4 | NA | -9999 | 107 | 349 | 105 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 10 | NA | -9999 | 186 | 441 | 186 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 5 | NA | -9999 | 134 | 352 | 133 | PASS | 0 | 40 | QC1 | Probe | 111 | 110 | 110 | PASS | QC2 | Probe | 126 | 126 | 126 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 21.1 | -9.2 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_05_0_02X | 614343860 | 92303 | Negative Control 1 | 847922 | 667585 | 6.8 | 2025-10-26T09:42:44+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 25.8 | 139 | PASS | -9999 | 176 | 368 | 172 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 6 | NA | -9999 | 171 | 404 | 169 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 1 | NA | -9999 | 238 | 584 | 237 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 2 | NA | -9999 | 96 | 310 | 94 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 7 | NA | -9999 | 148 | 340 | 148 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 2 | NA | -9999 | 135 | 355 | 133 | PASS | 0 | 40 | QC1 | Probe | 113 | 113 | 112 | PASS | QC2 | Probe | 98 | 98 | 98 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 20.1 | -8.4 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_05_0_03X | 614343496 | 92303 | Negative Control 1 | 847922 | 665815 | 6.8 | 2025-10-26T09:43:10+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 25 | 157 | PASS | -9999 | 215 | 444 | 209 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 4 | NA | -9999 | 187 | 442 | 186 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 8 | NA | -9999 | 258 | 639 | 258 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 2 | NA | -9999 | 97 | 300 | 96 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 4 | NA | -9999 | 181 | 431 | 181 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 7 | NA | -9999 | 116 | 316 | 114 | PASS | 0 | 40 | QC1 | Probe | 93 | 93 | 92 | PASS | QC2 | Probe | 105 | 105 | 104 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 18.9 | -7.9 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_05_0_04X | 614342592 | 92303 | Negative Control 1 | 847922 | 633201 | 6.8 | 2025-10-26T09:43:34+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 24.4 | 150 | PASS | -9999 | 182 | 383 | 176 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 5 | NA | -9999 | 154 | 372 | 152 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 8 | NA | -9999 | 207 | 523 | 208 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 4 | NA | -9999 | 89 | 277 | 92 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 5 | NA | -9999 | 159 | 357 | 162 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 2 | NA | -9999 | 116 | 315 | 117 | PASS | 0 | 40 | QC1 | Probe | 100 | 99 | 99 | PASS | QC2 | Probe | 120 | 119 | 119 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 21.8 | -7.2 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_05_0_05X | 614342822 | 92303 | Negative Control 1 | 847922 | 865593 | 6.8 | 2025-10-26T09:43:58+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 25.9 | 146 | PASS | -9999 | 200 | 435 | 196 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 12 | NA | -9999 | 186 | 443 | 183 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 11 | NA | -9999 | 250 | 618 | 247 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 6 | NA | -9999 | 108 | 327 | 106 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 6 | NA | -9999 | 195 | 466 | 194 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 6 | NA | -9999 | 146 | 377 | 144 | PASS | 0 | 40 | QC1 | Probe | 120 | 120 | 120 | PASS | QC2 | Probe | 109 | 109 | 109 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 22.3 | -12.8 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_05_0_06+ | 614343315 | 92303 | Negative Control 1 | 847922 | 791277 | 6.8 | 2025-10-26T09:44:22+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 24.1 | 157 | PASS | -9999 | 174 | 371 | 168 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 5 | NA | -9999 | 179 | 426 | 176 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 3 | NA | -9999 | 256 | 635 | 254 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 3 | NA | -9999 | 97 | 291 | 95 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 7 | NA | -9999 | 141 | 326 | 140 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 5 | NA | -9999 | 114 | 295 | 112 | PASS | 0 | 40 | QC1 | Probe | 100 | 100 | 100 | PASS | QC2 | Probe | 98 | 99 | 98 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 25.8 | -13.7 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_05_0_07+ | 614343631 | 92303 | Negative Control 1 | 847922 | 636287 | 6.8 | 2025-10-26T09:44:46+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 23.9 | 147 | PASS | -9999 | 166 | 343 | 162 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 2 | NA | -9999 | 162 | 380 | 160 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 0 | NA | -9999 | 223 | 553 | 220 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 0 | NA | -9999 | 84 | 252 | 82 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 6 | NA | -9999 | 131 | 303 | 130 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | -2 | NA | -9999 | 101 | 270 | 99 | PASS | 0 | 40 | QC1 | Probe | 83 | 83 | 83 | PASS | QC2 | Probe | 88 | 88 | 88 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 25.7 | -13.8 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_05_0_08+ | 614342981 | 92303 | Negative Control 1 | 847922 | 730163 | 6.8 | 2025-10-26T09:45:11+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 24.7 | 140 | PASS | -9999 | 189 | 386 | 184 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 4 | NA | -9999 | 177 | 422 | 176 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | -8 | NA | -9999 | 243 | 604 | 241 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 3 | NA | -9999 | 93 | 287 | 91 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 0 | NA | -9999 | 149 | 347 | 148 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 4 | NA | -9999 | 115 | 311 | 114 | PASS | 0 | 40 | QC1 | Probe | 95 | 95 | 95 | PASS | QC2 | Probe | 100 | 100 | 100 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 20 | -10.5 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_05_0_09+ | 614342683 | 92303 | Negative Control 1 | 847922 | 670395 | 6.8 | 2025-10-26T09:45:36+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 24.8 | 140 | PASS | -9999 | 147 | 305 | 143 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 5 | NA | -9999 | 163 | 384 | 161 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 7 | NA | -9999 | 220 | 546 | 218 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 3 | NA | -9999 | 95 | 297 | 94 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 6 | NA | -9999 | 140 | 311 | 139 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 4 | NA | -9999 | 140 | 360 | 138 | PASS | 0 | 40 | QC1 | Probe | 131 | 130 | 130 | PASS | QC2 | Probe | 93 | 92 | 92 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 16.2 | -7.3 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_05_0_10+ | 614342383 | 92303 | Negative Control 1 | 847922 | 210047638 | 6.8 | 2025-10-26T09:46:00+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 24.8 | 150 | PASS | -9999 | 227 | 489 | 221 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 5 | NA | -9999 | 209 | 498 | 207 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 6 | NA | -9999 | 270 | 672 | 268 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 3 | NA | -9999 | 121 | 365 | 119 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 5 | NA | -9999 | 160 | 388 | 160 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 6 | NA | -9999 | 156 | 405 | 154 | PASS | 0 | 40 | QC1 | Probe | 129 | 129 | 128 | PASS | QC2 | Probe | 110 | 110 | 110 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 17.9 | -9 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_06_0_01X | 614346890 | 92303 | Negative Control 1 | 847922 | 886194 | 6.8 | 2025-10-26T14:28:54+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 25 | 146 | PASS | -9999 | 214 | 466 | 200 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 10 | NA | -9999 | 182 | 434 | 177 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 7 | NA | -9999 | 243 | 595 | 230 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 2 | NA | -9999 | 110 | 339 | 109 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | -1 | NA | -9999 | 194 | 454 | 195 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 8 | NA | -9999 | 138 | 357 | 135 | PASS | 0 | 40 | QC1 | Probe | 130 | 129 | 128 | PASS | QC2 | Probe | 128 | 128 | 127 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 18.3 | -6.7 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_06_0_02X | 614347344 | 92303 | Negative Control 1 | 847922 | 797004 | 6.8 | 2025-10-26T14:29:19+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 25.8 | 149 | PASS | -9999 | 233 | 517 | 226 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 4 | NA | -9999 | 211 | 501 | 208 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 7 | NA | -9999 | 267 | 660 | 263 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 1 | NA | -9999 | 113 | 342 | 111 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 5 | NA | -9999 | 182 | 441 | 180 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 4 | NA | -9999 | 147 | 383 | 144 | PASS | 0 | 40 | QC1 | Probe | 118 | 117 | 117 | PASS | QC2 | Probe | 110 | 110 | 110 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 28.2 | -15.4 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_06_0_03X | 614347612 | 92303 | Negative Control 1 | 847922 | 742313 | 6.8 | 2025-10-26T14:29:43+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 24 | 154 | PASS | -9999 | 220 | 468 | 210 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 5 | NA | -9999 | 210 | 497 | 205 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 10 | NA | -9999 | 256 | 619 | 250 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 0 | NA | -9999 | 112 | 344 | 110 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 7 | NA | -9999 | 160 | 382 | 158 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 6 | NA | -9999 | 149 | 401 | 145 | PASS | 0 | 40 | QC1 | Probe | 112 | 111 | 111 | PASS | QC2 | Probe | 109 | 109 | 109 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 22.9 | -10.5 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_06_0_04X | 614345900 | 92303 | Negative Control 1 | 847922 | 604077 | 6.8 | 2025-10-26T14:30:07+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 24.5 | 146 | PASS | -9999 | 203 | 450 | 200 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 7 | NA | -9999 | 213 | 501 | 211 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 8 | NA | -9999 | 247 | 606 | 245 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 1 | NA | -9999 | 109 | 326 | 107 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 3 | NA | -9999 | 173 | 423 | 173 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 6 | NA | -9999 | 151 | 396 | 148 | PASS | 0 | 40 | QC1 | Probe | 132 | 132 | 131 | PASS | QC2 | Probe | 119 | 118 | 118 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 17.7 | -7.4 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_06_0_05X | 614346594 | 92303 | Negative Control 1 | 847922 | 617583 | 6.8 | 2025-10-26T14:30:32+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 23.7 | 137 | PASS | -9999 | 211 | 448 | 201 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 7 | NA | -9999 | 200 | 473 | 194 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 9 | NA | -9999 | 270 | 662 | 264 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 2 | NA | -9999 | 91 | 282 | 88 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 5 | NA | -9999 | 172 | 401 | 170 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 7 | NA | -9999 | 147 | 389 | 143 | PASS | 0 | 40 | QC1 | Probe | 132 | 132 | 132 | PASS | QC2 | Probe | 117 | 117 | 116 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 17.5 | -7.9 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_06_0_06+ | 614347035 | 92303 | Negative Control 1 | 847922 | 725382 | 6.8 | 2025-10-26T14:30:56+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 25.9 | 135 | PASS | -9999 | 208 | 459 | 204 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 5 | NA | -9999 | 217 | 515 | 215 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 6 | NA | -9999 | 288 | 709 | 285 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 1 | NA | -9999 | 100 | 308 | 98 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 8 | NA | -9999 | 172 | 396 | 170 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 2 | NA | -9999 | 141 | 383 | 139 | PASS | 0 | 40 | QC1 | Probe | 121 | 121 | 120 | PASS | QC2 | Probe | 113 | 112 | 112 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 28.3 | -18.7 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_06_0_07+ | 614346097 | 92303 | Negative Control 1 | 847922 | 710982 | 6.8 | 2025-10-26T14:31:22+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 24.2 | 142 | PASS | -9999 | 205 | 447 | 200 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 7 | NA | -9999 | 181 | 435 | 179 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | -2 | NA | -9999 | 241 | 595 | 241 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 3 | NA | -9999 | 96 | 307 | 95 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 0 | NA | -9999 | 191 | 447 | 192 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 4 | NA | -9999 | 137 | 358 | 135 | PASS | 0 | 40 | QC1 | Probe | 118 | 117 | 117 | PASS | QC2 | Probe | 131 | 131 | 131 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 19.7 | -7.2 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_06_0_08+ | 614347803 | 92303 | Negative Control 1 | 847922 | 657194 | 6.8 | 2025-10-26T14:31:46+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 24.1 | 158 | PASS | -9999 | 219 | 481 | 210 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 3 | NA | -9999 | 193 | 456 | 190 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 2 | NA | -9999 | 261 | 649 | 259 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 0 | NA | -9999 | 109 | 327 | 106 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 8 | NA | -9999 | 153 | 372 | 154 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 3 | NA | -9999 | 148 | 384 | 145 | PASS | 0 | 40 | QC1 | Probe | 115 | 115 | 114 | PASS | QC2 | Probe | 105 | 105 | 105 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 24.1 | -12.7 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_06_0_09+ | 614346729 | 92303 | Negative Control 1 | 847922 | 702829 | 6.8 | 2025-10-26T14:32:09+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 24 | 145 | PASS | -9999 | 221 | 489 | 216 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 6 | NA | -9999 | 193 | 464 | 192 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 8 | NA | -9999 | 247 | 611 | 246 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 0 | NA | -9999 | 114 | 350 | 113 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 3 | NA | -9999 | 174 | 417 | 174 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 4 | NA | -9999 | 153 | 401 | 152 | PASS | 0 | 40 | QC1 | Probe | 131 | 131 | 131 | PASS | QC2 | Probe | 125 | 125 | 125 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 18.8 | -9 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_06_0_10+ | 614347201 | 92303 | Negative Control 1 | 847922 | 220000161 | 6.8 | 2025-10-26T14:32:31+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 23.7 | 156 | PASS | -9999 | 215 | 460 | 203 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 7 | NA | -9999 | 198 | 471 | 195 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 4 | NA | -9999 | 261 | 638 | 256 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 4 | NA | -9999 | 110 | 328 | 107 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 13 | NA | -9999 | 148 | 359 | 147 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 8 | NA | -9999 | 147 | 383 | 144 | PASS | 0 | 40 | QC1 | Probe | 115 | 114 | 113 | PASS | QC2 | Probe | 104 | 103 | 103 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 21.2 | -9.8 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_07_0_01X | 614351832 | 92303 | Negative Control 1 | 847922 | 886194 | 6.8 | 2025-10-26T15:58:21+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 24.5 | 149 | PASS | -9999 | 234 | 509 | 225 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 6 | NA | -9999 | 198 | 475 | 196 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 6 | NA | -9999 | 253 | 634 | 250 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 2 | NA | -9999 | 117 | 357 | 116 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 2 | NA | -9999 | 201 | 477 | 201 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 5 | NA | -9999 | 142 | 371 | 140 | PASS | 0 | 40 | QC1 | Probe | 135 | 133 | 132 | PASS | QC2 | Probe | 143 | 143 | 143 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 29.5 | -18 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_07_0_02X | 614351590 | 92303 | Negative Control 1 | 847922 | 797004 | 6.8 | 2025-10-26T15:58:46+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 24.1 | 154 | PASS | -9999 | 231 | 503 | 224 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 7 | NA | -9999 | 209 | 496 | 205 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 14 | NA | -9999 | 260 | 639 | 257 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 2 | NA | -9999 | 111 | 336 | 110 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 4 | NA | -9999 | 180 | 429 | 179 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 6 | NA | -9999 | 149 | 390 | 146 | PASS | 0 | 40 | QC1 | Probe | 123 | 123 | 122 | PASS | QC2 | Probe | 113 | 113 | 113 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 27.6 | -14 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_07_0_03X | 614351220 | 92303 | Negative Control 1 | 847922 | 742313 | 6.8 | 2025-10-26T15:59:10+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 24.7 | 154 | PASS | -9999 | 219 | 470 | 210 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 7 | NA | -9999 | 203 | 483 | 199 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 9 | NA | -9999 | 255 | 626 | 251 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 0 | NA | -9999 | 111 | 342 | 109 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 8 | NA | -9999 | 167 | 388 | 166 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 5 | NA | -9999 | 139 | 375 | 135 | PASS | 0 | 40 | QC1 | Probe | 106 | 106 | 106 | PASS | QC2 | Probe | 101 | 101 | 101 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 25.6 | -11.9 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_07_0_04X | 614350974 | 92303 | Negative Control 1 | 847922 | 604077 | 6.8 | 2025-10-26T15:59:32+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 25.7 | 146 | PASS | -9999 | 197 | 436 | 192 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 3 | NA | -9999 | 205 | 484 | 201 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 8 | NA | -9999 | 237 | 575 | 233 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 4 | NA | -9999 | 107 | 321 | 106 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 2 | NA | -9999 | 167 | 406 | 166 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 6 | NA | -9999 | 127 | 334 | 125 | PASS | 0 | 40 | QC1 | Probe | 111 | 111 | 110 | PASS | QC2 | Probe | 103 | 102 | 102 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 23.7 | -7.5 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_07_0_05X | 614350660 | 92303 | Negative Control 1 | 847922 | 617583 | 6.8 | 2025-10-26T15:59:57+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 24.2 | 126 | PASS | -9999 | 213 | 451 | 203 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 4 | NA | -9999 | 197 | 470 | 191 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 3 | NA | -9999 | 265 | 652 | 259 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 0 | NA | -9999 | 91 | 279 | 89 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 3 | NA | -9999 | 167 | 391 | 165 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 3 | NA | -9999 | 142 | 381 | 138 | PASS | 0 | 40 | QC1 | Probe | 132 | 131 | 131 | PASS | QC2 | Probe | 117 | 117 | 117 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 23.9 | -18.2 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_07_0_06+ | 614350499 | 92303 | Negative Control 1 | 847922 | 725382 | 6.8 | 2025-10-26T16:00:22+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 23.8 | 147 | PASS | -9999 | 206 | 452 | 204 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 4 | NA | -9999 | 223 | 531 | 221 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 6 | NA | -9999 | 292 | 721 | 290 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 3 | NA | -9999 | 102 | 315 | 100 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 8 | NA | -9999 | 175 | 405 | 174 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 4 | NA | -9999 | 150 | 404 | 148 | PASS | 0 | 40 | QC1 | Probe | 130 | 129 | 129 | PASS | QC2 | Probe | 123 | 123 | 123 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 24.4 | -19.2 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_07_0_07+ | 614351129 | 92303 | Negative Control 1 | 847922 | 710982 | 6.8 | 2025-10-26T16:00:48+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 24.8 | 135 | PASS | -9999 | 215 | 478 | 209 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 8 | NA | -9999 | 184 | 439 | 182 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | -6 | NA | -9999 | 259 | 645 | 258 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 2 | NA | -9999 | 102 | 313 | 100 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 0 | NA | -9999 | 194 | 473 | 194 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 5 | NA | -9999 | 131 | 343 | 129 | PASS | 0 | 40 | QC1 | Probe | 116 | 115 | 115 | PASS | QC2 | Probe | 114 | 114 | 114 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 27 | -13 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_07_0_08+ | 614350803 | 92303 | Negative Control 1 | 847922 | 657194 | 6.8 | 2025-10-26T16:01:10+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 26.2 | 149 | PASS | -9999 | 238 | 528 | 231 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 4 | NA | -9999 | 206 | 489 | 204 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 2 | NA | -9999 | 287 | 718 | 286 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 2 | NA | -9999 | 116 | 350 | 115 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 7 | NA | -9999 | 165 | 410 | 165 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 3 | NA | -9999 | 183 | 475 | 181 | PASS | 0 | 40 | QC1 | Probe | 116 | 115 | 115 | PASS | QC2 | Probe | 97 | 96 | 96 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 24 | -12.9 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_07_0_09+ | 614351713 | 92303 | Negative Control 1 | 847922 | 702829 | 6.8 | 2025-10-26T16:01:34+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 23.9 | 152 | PASS | -9999 | 223 | 494 | 218 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 6 | NA | -9999 | 202 | 479 | 199 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 13 | NA | -9999 | 254 | 626 | 250 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 2 | NA | -9999 | 116 | 353 | 114 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 4 | NA | -9999 | 175 | 423 | 175 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 6 | NA | -9999 | 155 | 407 | 153 | PASS | 0 | 40 | QC1 | Probe | 137 | 137 | 136 | PASS | QC2 | Probe | 114 | 114 | 114 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 25.4 | -13.1 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_07_0_10+ | 614351397 | 92303 | Negative Control 1 | 847922 | 220000161 | 6.8 | 2025-10-26T16:01:56+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 24.1 | 156 | PASS | -9999 | 214 | 457 | 202 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 9 | NA | -9999 | 191 | 453 | 188 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 1 | NA | -9999 | 262 | 644 | 257 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 2 | NA | -9999 | 106 | 315 | 104 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 10 | NA | -9999 | 152 | 358 | 151 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 5 | NA | -9999 | 145 | 377 | 141 | PASS | 0 | 40 | QC1 | Probe | 113 | 113 | 112 | PASS | QC2 | Probe | 105 | 105 | 105 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 23.4 | -12.6 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_08_0_01X | 614383328 | 92303 | Negative Control 1 | 847922 | 886194 | 6.8 | 2025-10-26T18:28:41+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 24.8 | 126 | PASS | -9999 | 196 | 434 | 182 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 3 | NA | -9999 | 184 | 429 | 177 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | -9 | NA | -9999 | 239 | 585 | 233 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 0 | NA | -9999 | 109 | 323 | 104 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | -3 | NA | -9999 | 177 | 423 | 173 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 2 | NA | -9999 | 131 | 335 | 126 | PASS | 0 | 40 | QC1 | Probe | 113 | 112 | 111 | PASS | QC2 | Probe | 110 | 109 | 109 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 28.9 | -20.9 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_08_0_02X | 614383866 | 92303 | Negative Control 1 | 847922 | 797004 | 6.8 | 2025-10-26T18:29:31+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 23.6 | 149 | PASS | -9999 | 222 | 484 | 215 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 7 | NA | -9999 | 197 | 471 | 194 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 10 | NA | -9999 | 253 | 620 | 249 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 1 | NA | -9999 | 112 | 341 | 110 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 3 | NA | -9999 | 181 | 427 | 180 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 5 | NA | -9999 | 147 | 384 | 144 | PASS | 0 | 40 | QC1 | Probe | 123 | 123 | 122 | PASS | QC2 | Probe | 117 | 116 | 116 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 16.6 | -8.4 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_08_0_03X | 614352842 | 92303 | Negative Control 1 | 847922 | 742313 | 6.8 | 2025-10-26T18:30:44+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 25.5 | 151 | PASS | -9999 | 216 | 468 | 207 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 5 | NA | -9999 | 206 | 492 | 202 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 6 | NA | -9999 | 255 | 625 | 251 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | -2 | NA | -9999 | 112 | 348 | 110 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 9 | NA | -9999 | 171 | 390 | 169 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 3 | NA | -9999 | 137 | 374 | 134 | PASS | 0 | 40 | QC1 | Probe | 112 | 111 | 111 | PASS | QC2 | Probe | 113 | 112 | 112 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 21.6 | -10 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_08_0_04X | 614384298 | 92303 | Negative Control 1 | 847922 | 604077 | 6.8 | 2025-10-26T18:31:55+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 26.1 | 132 | PASS | -9999 | 184 | 411 | 180 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 6 | NA | -9999 | 194 | 458 | 191 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 1 | NA | -9999 | 223 | 541 | 221 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | -1 | NA | -9999 | 103 | 311 | 101 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 1 | NA | -9999 | 168 | 395 | 167 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 4 | NA | -9999 | 140 | 366 | 138 | PASS | 0 | 40 | QC1 | Probe | 114 | 113 | 113 | PASS | QC2 | Probe | 97 | 97 | 97 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 23.1 | -8.4 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_08_0_05X | 614352910 | 92303 | Negative Control 1 | 847922 | 617583 | 6.8 | 2025-10-26T18:33:10+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 23.4 | 142 | PASS | -9999 | 208 | 440 | 195 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 8 | NA | -9999 | 192 | 457 | 184 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 12 | NA | -9999 | 260 | 646 | 253 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 2 | NA | -9999 | 93 | 281 | 90 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 6 | NA | -9999 | 170 | 395 | 167 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 9 | NA | -9999 | 142 | 375 | 136 | PASS | 0 | 40 | QC1 | Probe | 133 | 133 | 132 | PASS | QC2 | Probe | 110 | 109 | 109 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 19.4 | -9 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_08_0_06+ | 614383509 | 92303 | Negative Control 1 | 847922 | 725382 | 6.8 | 2025-10-26T18:33:34+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 24.7 | 133 | PASS | -9999 | 193 | 418 | 191 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 6 | NA | -9999 | 199 | 476 | 198 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 6 | NA | -9999 | 265 | 652 | 265 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 1 | NA | -9999 | 97 | 300 | 96 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 10 | NA | -9999 | 166 | 361 | 165 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 1 | NA | -9999 | 147 | 396 | 146 | PASS | 0 | 40 | QC1 | Probe | 132 | 131 | 131 | PASS | QC2 | Probe | 119 | 118 | 118 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 25.2 | -12.7 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_08_0_07+ | 614382571 | 92303 | Negative Control 1 | 847922 | 710982 | 6.8 | 2025-10-26T18:34:00+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 24.1 | 136 | PASS | -9999 | 204 | 449 | 200 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 7 | NA | -9999 | 179 | 430 | 176 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | -4 | NA | -9999 | 240 | 599 | 241 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 2 | NA | -9999 | 98 | 305 | 96 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 1 | NA | -9999 | 186 | 446 | 187 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 5 | NA | -9999 | 130 | 342 | 129 | PASS | 0 | 40 | QC1 | Probe | 114 | 113 | 113 | PASS | QC2 | Probe | 117 | 117 | 117 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 28.7 | -14 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_08_0_08+ | 614383987 | 92303 | Negative Control 1 | 847922 | 912152 | 6.8 | 2025-10-26T18:34:23+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 23.2 | 163 | PASS | -9999 | 195 | 434 | 187 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 6 | NA | -9999 | 218 | 511 | 214 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 6 | NA | -9999 | 266 | 653 | 261 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 7 | NA | -9999 | 104 | 334 | 101 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 7 | NA | -9999 | 172 | 417 | 171 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 7 | NA | -9999 | 147 | 390 | 144 | PASS | 0 | 40 | QC1 | Probe | 121 | 120 | 119 | PASS | QC2 | Probe | 129 | 129 | 128 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 16.7 | -6.8 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_08_0_09+ | 614384411 | 92303 | Negative Control 1 | 847922 | 784559 | 6.8 | 2025-10-26T18:34:48+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 23.5 | 133 | PASS | -9999 | 204 | 414 | 201 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 4 | NA | -9999 | 188 | 448 | 187 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 7 | NA | -9999 | 259 | 625 | 259 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 4 | NA | -9999 | 113 | 347 | 112 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 4 | NA | -9999 | 182 | 420 | 182 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 4 | NA | -9999 | 187 | 484 | 185 | PASS | 0 | 40 | QC1 | Probe | 136 | 135 | 135 | PASS | QC2 | Probe | 135 | 135 | 134 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 17.4 | -7 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_08_0_10+ | 614352459 | 92303 | Negative Control 1 | 847922 | 210053180 | 6.8 | 2025-10-26T18:35:13+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 23.8 | 154 | PASS | -9999 | 226 | 497 | 220 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 7 | NA | -9999 | 212 | 506 | 209 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 14 | NA | -9999 | 288 | 709 | 285 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 4 | NA | -9999 | 117 | 369 | 115 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 4 | NA | -9999 | 211 | 496 | 211 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 7 | NA | -9999 | 181 | 470 | 179 | PASS | 0 | 40 | QC1 | Probe | 150 | 150 | 149 | PASS | QC2 | Probe | 130 | 129 | 129 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 29.2 | -22.7 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_09_0_01X | 614387708 | 92303 | Negative Control 1 | 847922 | 686063 | 6.8 | 2025-10-26T21:41:21+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 24.2 | 129 | PASS | -9999 | 141 | 312 | 134 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | -2 | NA | -9999 | 143 | 340 | 140 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | -11 | NA | -9999 | 206 | 504 | 201 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 0 | NA | -9999 | 92 | 274 | 90 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 6 | NA | -9999 | 123 | 268 | 120 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 2 | NA | -9999 | 109 | 281 | 106 | PASS | 0 | 40 | QC1 | Probe | 100 | 99 | 99 | PASS | QC2 | Probe | 92 | 92 | 91 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 18.1 | -4.9 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_09_0_02X | 614388006 | 92303 | Negative Control 1 | 847922 | 684250 | 6.8 | 2025-10-26T21:41:45+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 23.9 | 150 | PASS | -9999 | 195 | 426 | 190 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 6 | NA | -9999 | 167 | 400 | 165 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 4 | NA | -9999 | 232 | 572 | 231 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 3 | NA | -9999 | 104 | 321 | 103 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 6 | NA | -9999 | 186 | 439 | 186 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 9 | NA | -9999 | 134 | 350 | 132 | PASS | 0 | 40 | QC1 | Probe | 115 | 115 | 114 | PASS | QC2 | Probe | 116 | 116 | 116 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 19.8 | -9 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_09_0_03X | 614387446 | 92303 | Negative Control 1 | 847922 | 790976 | 6.8 | 2025-10-26T21:42:10+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 25 | 138 | PASS | -9999 | 178 | 391 | 179 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 1 | NA | -9999 | 196 | 467 | 195 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 2 | NA | -9999 | 271 | 673 | 269 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 1 | NA | -9999 | 116 | 350 | 114 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 4 | NA | -9999 | 178 | 423 | 177 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 6 | NA | -9999 | 140 | 369 | 138 | PASS | 0 | 40 | QC1 | Probe | 116 | 115 | 115 | PASS | QC2 | Probe | 110 | 111 | 110 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 16.6 | -7.7 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_09_0_04X | 614387172 | 92303 | Negative Control 1 | 847922 | 654187 | 6.8 | 2025-10-26T21:42:34+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 26 | 135 | PASS | -9999 | 204 | 423 | 202 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 11 | NA | -9999 | 159 | 379 | 158 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 23 | NA | -9999 | 241 | 578 | 238 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 10 | NA | -9999 | 98 | 307 | 96 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 9 | NA | -9999 | 184 | 404 | 183 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 10 | NA | -9999 | 123 | 312 | 121 | PASS | 0 | 40 | QC1 | Probe | 106 | 106 | 106 | PASS | QC2 | Probe | 112 | 112 | 112 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 19.2 | -9.2 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_09_0_05X | 614386812 | 92303 | Negative Control 1 | 847922 | 674666 | 6.8 | 2025-10-26T21:42:57+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 25 | 148 | PASS | -9999 | 194 | 430 | 188 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 1 | NA | -9999 | 213 | 508 | 210 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 4 | NA | -9999 | 264 | 654 | 260 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 1 | NA | -9999 | 102 | 313 | 100 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 3 | NA | -9999 | 166 | 399 | 165 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 3 | NA | -9999 | 122 | 316 | 119 | PASS | 0 | 40 | QC1 | Probe | 99 | 99 | 98 | PASS | QC2 | Probe | 113 | 113 | 113 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 19 | -9.9 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_09_0_06+ | 614386569 | 92303 | Negative Control 1 | 847922 | 797456 | 6.8 | 2025-10-26T21:43:20+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 25 | 125 | PASS | -9999 | 178 | 387 | 172 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 5 | NA | -9999 | 166 | 398 | 165 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 1 | NA | -9999 | 223 | 551 | 222 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 2 | NA | -9999 | 91 | 277 | 89 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 1 | NA | -9999 | 163 | 383 | 163 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 7 | NA | -9999 | 119 | 308 | 118 | PASS | 0 | 40 | QC1 | Probe | 98 | 98 | 98 | PASS | QC2 | Probe | 95 | 95 | 95 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 18.1 | -9.2 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_09_0_07+ | 614387311 | 92303 | Negative Control 1 | 847922 | 695878 | 6.8 | 2025-10-26T21:43:46+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 23.8 | 155 | PASS | -9999 | 215 | 467 | 211 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 6 | NA | -9999 | 193 | 463 | 193 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 6 | NA | -9999 | 287 | 713 | 286 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 4 | NA | -9999 | 109 | 332 | 110 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 5 | NA | -9999 | 181 | 418 | 181 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 6 | NA | -9999 | 130 | 342 | 129 | PASS | 0 | 40 | QC1 | Probe | 118 | 117 | 117 | PASS | QC2 | Probe | 121 | 121 | 121 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 21.8 | -7.9 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_09_0_08+ | 614387595 | 92303 | Negative Control 1 | 847922 | 784143 | 6.8 | 2025-10-26T21:44:12+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 24.1 | 149 | PASS | -9999 | 222 | 489 | 214 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | -1 | NA | -9999 | 191 | 452 | 189 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 3 | NA | -9999 | 258 | 635 | 255 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | -1 | NA | -9999 | 117 | 352 | 115 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 5 | NA | -9999 | 184 | 424 | 182 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 5 | NA | -9999 | 138 | 360 | 136 | PASS | 0 | 40 | QC1 | Probe | 122 | 121 | 120 | PASS | QC2 | Probe | 109 | 109 | 109 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 18.9 | -9 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_09_0_09+ | 614387921 | 92303 | Negative Control 1 | 847922 | 673187 | 6.8 | 2025-10-26T21:44:35+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 24.4 | 168 | PASS | -9999 | 211 | 467 | 207 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 11 | NA | -9999 | 213 | 506 | 210 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 15 | NA | -9999 | 266 | 664 | 261 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 2 | NA | -9999 | 102 | 303 | 101 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 8 | NA | -9999 | 187 | 425 | 185 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 6 | NA | -9999 | 108 | 284 | 106 | PASS | 0 | 40 | QC1 | Probe | 107 | 106 | 106 | PASS | QC2 | Probe | 107 | 106 | 106 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 21.9 | -10.4 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_09_0_10+ | 614386955 | 92303 | Negative Control 1 | 847922 | 888731 | 6.8 | 2025-10-26T21:44:59+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 23.9 | 146 | PASS | -9999 | 186 | 425 | 188 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 11 | NA | -9999 | 155 | 407 | 167 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 10 | NA | -9999 | 230 | 606 | 240 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 1 | NA | -9999 | 98 | 313 | 104 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 7 | NA | -9999 | 143 | 358 | 154 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 5 | NA | -9999 | 124 | 348 | 131 | PASS | 0 | 40 | QC1 | Probe | 111 | 110 | 110 | PASS | QC2 | Probe | 106 | 105 | 105 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 18.5 | -8.2 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_10_0_01X | 614390504 | 92303 | Negative Control 1 | 847922 | 657194 | 6.8 | 2025-10-26T23:37:12+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 23.8 | 144 | PASS | -9999 | 205 | 434 | 195 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 3 | NA | -9999 | 179 | 424 | 176 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | -3 | NA | -9999 | 243 | 596 | 239 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 3 | NA | -9999 | 102 | 313 | 100 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 6 | NA | -9999 | 146 | 344 | 146 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 0 | NA | -9999 | 150 | 392 | 148 | PASS | 0 | 40 | QC1 | Probe | 115 | 115 | 115 | PASS | QC2 | Probe | 101 | 101 | 101 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 20.5 | -11.5 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_10_0_02X | 614390490 | 92303 | Negative Control 1 | 847922 | 702829 | 6.8 | 2025-10-26T23:38:00+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 24.6 | 145 | PASS | -9999 | 224 | 481 | 217 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 6 | NA | -9999 | 194 | 462 | 192 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 7 | NA | -9999 | 248 | 603 | 245 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 2 | NA | -9999 | 112 | 343 | 111 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 5 | NA | -9999 | 172 | 409 | 171 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 5 | NA | -9999 | 164 | 428 | 162 | PASS | 0 | 40 | QC1 | Probe | 124 | 124 | 124 | PASS | QC2 | Probe | 118 | 118 | 117 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 16.4 | -7.1 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_10_0_03X | 614390034 | 92303 | Negative Control 1 | 847922 | 220000161 | 6.8 | 2025-10-26T23:39:12+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 25.1 | 139 | PASS | -9999 | 237 | 519 | 228 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 7 | NA | -9999 | 202 | 479 | 200 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 2 | NA | -9999 | 285 | 707 | 282 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 0 | NA | -9999 | 112 | 331 | 109 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 4 | NA | -9999 | 158 | 388 | 157 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 3 | NA | -9999 | 145 | 376 | 142 | PASS | 0 | 40 | QC1 | Probe | 115 | 114 | 114 | PASS | QC2 | Probe | 95 | 95 | 95 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 20.9 | -9.3 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_10_0_04X | 614390494 | 92303 | Negative Control 1 | 847922 | 912152 | 6.8 | 2025-10-26T23:40:24+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 24.4 | 134 | PASS | -9999 | 203 | 445 | 198 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 4 | NA | -9999 | 207 | 494 | 205 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | -1 | NA | -9999 | 261 | 647 | 261 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 3 | NA | -9999 | 108 | 341 | 106 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 2 | NA | -9999 | 176 | 422 | 176 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 2 | NA | -9999 | 148 | 396 | 145 | PASS | 0 | 40 | QC1 | Probe | 123 | 122 | 122 | PASS | QC2 | Probe | 124 | 124 | 124 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 17.1 | -9 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_10_0_05X | 614390402 | 92303 | Negative Control 1 | 847922 | 210046047 | 6.8 | 2025-10-26T23:41:37+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 23.9 | 140 | PASS | -9999 | 238 | 525 | 228 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 7 | NA | -9999 | 190 | 456 | 187 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 0 | NA | -9999 | 257 | 641 | 256 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 0 | NA | -9999 | 99 | 302 | 97 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 5 | NA | -9999 | 179 | 424 | 179 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 5 | NA | -9999 | 139 | 367 | 137 | PASS | 0 | 40 | QC1 | Probe | 133 | 133 | 132 | PASS | QC2 | Probe | 120 | 120 | 120 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 20.5 | -10.1 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_10_0_06+ | 614390505 | 92303 | Negative Control 1 | 847922 | 701892 | 6.8 | 2025-10-26T23:42:50+0100 | Aborted | ERROR | SPC | SPC | 1,2,3,4 | PASS | 23.5 | 149 | PASS | -9999 | 231 | 507 | 224 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 1 | NA | -9999 | 226 | 531 | 223 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 6 | NA | -9999 | 282 | 697 | 278 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 3 | NA | -9999 | 117 | 350 | 115 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 7 | NA | -9999 | 204 | 485 | 202 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 6 | NA | -9999 | 159 | 415 | 155 | PASS | 0 | 40 | QC1 | Probe | 142 | 141 | 140 | PASS | QC2 | Probe | 151 | 151 | 150 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 92 | -8.2 | Error 2008: Syringe pressure reading of 100.1 PSI exceeds the protocol limit of 100.0 PSI |  |  |  |  | <None>
Xpert MTB-RIF Ultra | 4 | NEG_10_0_07+ | 614390503 | 92303 | Negative Control 1 | 847922 | 866331 | 6.8 | 2025-10-26T23:43:14+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 23.7 | 148 | PASS | -9999 | 212 | 454 | 200 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 12 | NA | -9999 | 181 | 432 | 176 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 14 | NA | -9999 | 218 | 526 | 214 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 1 | NA | -9999 | 95 | 298 | 91 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 7 | NA | -9999 | 175 | 402 | 172 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 10 | NA | -9999 | 136 | 355 | 131 | PASS | 0 | 40 | QC1 | Probe | 117 | 116 | 116 | PASS | QC2 | Probe | 105 | 105 | 105 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 13.4 | -6.3 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_10_0_08+ | 614389891 | 92303 | Negative Control 1 | 847922 | 902140 | 6.8 | 2025-10-26T23:43:37+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 26.4 | 143 | PASS | -9999 | 185 | 398 | 179 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 5 | NA | -9999 | 186 | 439 | 183 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 0 | NA | -9999 | 235 | 581 | 232 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 3 | NA | -9999 | 89 | 268 | 87 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 8 | NA | -9999 | 148 | 352 | 147 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 2 | NA | -9999 | 146 | 382 | 143 | PASS | 0 | 40 | QC1 | Probe | 118 | 117 | 117 | PASS | QC2 | Probe | 93 | 93 | 93 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 17.1 | -7.5 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_10_0_09+ | 614390529 | 92303 | Negative Control 1 | 847922 | 635023 | 6.8 | 2025-10-26T23:44:01+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 22.8 | 156 | PASS | -9999 | 169 | 373 | 164 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 8 | NA | -9999 | 197 | 472 | 195 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 10 | NA | -9999 | 259 | 634 | 258 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 6 | NA | -9999 | 112 | 343 | 111 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 10 | NA | -9999 | 173 | 387 | 171 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 10 | NA | -9999 | 132 | 349 | 130 | PASS | 0 | 40 | QC1 | Probe | 127 | 126 | 126 | PASS | QC2 | Probe | 117 | 116 | 116 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 18.1 | -9.8 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | NEG_10_0_10+ | 614390273 | 92303 | Negative Control 1 | 847922 | 652001 | 6.8 | 2025-10-26T23:44:25+0100 | Done | MTB NOT DETECTED | SPC | SPC | 1,2,3,4 | PASS | 23.9 | 162 | PASS | -9999 | 170 | 371 | 162 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | FAIL | 0 | 6 | NA | -9999 | 168 | 401 | 166 | PASS | 0 | 40 | rpoB1 | Target | 1 | INVALID | 0 | 10 | NA | -9999 | 227 | 559 | 224 | PASS | 0 | 40 | rpoB2 | Target | 2 | INVALID | 0 | 2 | NA | -9999 | 97 | 299 | 95 | PASS | 0 | 40 | rpoB3 | Target | 3 | INVALID | 0 | 7 | NA | -9999 | 145 | 346 | 145 | PASS | 0 | 40 | rpoB4 | Target | 4 | INVALID | 0 | 8 | NA | -9999 | 125 | 333 | 123 | PASS | 0 | 40 | QC1 | Probe | 108 | 107 | 107 | PASS | QC2 | Probe | 93 | 93 | 92 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 18.2 | -8.3 | [5 empty] | <None>
Enter Resample Data Below: |  | ¨N/A
Assay | Assay Version | Sample ID | Cartridge S/N | Reagent Lot ID | Test Type | Instrument S/N | Module S/N | S/W Version | Start Time | Status | Test Result | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Max Pressure (PSI) | Min Pressure (PSI) | Error | Patient ID | Quantitative Value | Quantitative Unit | Interpretation Flag | User

## Data - Sample Type 1
Assay | Assay Version | Sample ID | Cartridge S/N | Reagent Lot ID | Test Type | Instrument S/N | Module S/N | S/W Version | Start Time | Status | Test Result | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Max Pressure (PSI) | Min Pressure (PSI) | Error | Patient ID | Quantitative Value | Quantitative Unit | Interpretation Flag | User
Xpert MTB-RIF Ultra | 4 | C15NOV23D_01_1_11X | 614332200 | 92303 | Positive Control 1 | 839032 | 652304 | 6.8 | 2025-10-26T01:52:14+0200 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 25.3 | 140 | NA | -9999 | 226 | 495 | 221 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 27.2 | 117 | NA | -9999 | 209 | 504 | 208 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 25 | 382 | PASS | -9999 | 274 | 683 | 272 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 25.9 | 172 | PASS | -9999 | 119 | 381 | 117 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 30 | NA | -9999 | 189 | 457 | 189 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 23.5 | 396 | PASS | -9999 | 137 | 359 | 135 | PASS | 0 | 40 | QC1 | Probe | 113 | 112 | 112 | PASS | QC2 | Probe | 111 | 111 | 111 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.3 | 50.5 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70 | 46.6 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.5 | 148.9 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 21.2 | -19.7 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_01_1_12X | 614332198 | 92303 | Positive Control 1 | 839032 | 653259 | 6.8 | 2025-10-26T01:52:39+0200 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 24.1 | 166 | NA | -9999 | 231 | 496 | 222 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 25.6 | 127 | NA | -9999 | 206 | 491 | 202 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 23.5 | 433 | PASS | -9999 | 250 | 616 | 246 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24.8 | 180 | PASS | -9999 | 123 | 371 | 122 | PASS | 0 | 40 | rpoB3 | Target | 3 | POS | 39.2 | 41 | PASS | -9999 | 182 | 441 | 181 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 22.3 | 407 | PASS | -9999 | 154 | 401 | 151 | PASS | 0 | 40 | QC1 | Probe | 121 | 120 | 120 | PASS | QC2 | Probe | 114 | 114 | 114 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.2 | 46.4 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70 | 49.3 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.5 | 175.3 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 19.8 | -11.3 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_01_1_13X | 614332204 | 92303 | Positive Control 1 | 839032 | 627453 | 6.8 | 2025-10-26T01:53:06+0200 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 24.1 | 161 | NA | -9999 | 258 | 545 | 254 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 26.3 | 130 | NA | -9999 | 206 | 486 | 203 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 23.5 | 433 | PASS | -9999 | 279 | 688 | 277 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24.6 | 186 | PASS | -9999 | 104 | 309 | 102 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 37 | NA | -9999 | 174 | 421 | 173 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 22.2 | 434 | PASS | -9999 | 128 | 331 | 126 | PASS | 0 | 40 | QC1 | Probe | 106 | 106 | 106 | PASS | QC2 | Probe | 109 | 109 | 108 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.1 | 49.1 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 69.9 | 43.6 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.4 | 149.1 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 18.4 | -12.9 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_01_1_14X | 614332202 | 92303 | Positive Control 1 | 839032 | 782120 | 6.8 | 2025-10-26T01:53:32+0200 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 24.8 | 147 | NA | -9999 | 213 | 431 | 208 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 26.4 | 115 | NA | -9999 | 209 | 492 | 207 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 23.8 | 411 | PASS | -9999 | 298 | 726 | 295 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24.9 | 178 | PASS | -9999 | 109 | 333 | 107 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 34 | NA | -9999 | 179 | 420 | 178 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 22.6 | 403 | PASS | -9999 | 149 | 397 | 147 | PASS | 0 | 40 | QC1 | Probe | 115 | 114 | 114 | PASS | QC2 | Probe | 113 | 113 | 113 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.1 | 66.4 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 69.8 | 50.9 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.3 | 169.8 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 24.8 | -16.8 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_01_1_15+ | 614332201 | 92303 | Positive Control 1 | 839032 | 210034622 | 6.8 | 2025-10-26T01:53:57+0200 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 23.9 | 156 | NA | -9999 | 169 | 375 | 164 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 26.8 | 116 | NA | -9999 | 173 | 416 | 171 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 24 | 396 | PASS | -9999 | 234 | 581 | 233 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24.7 | 182 | PASS | -9999 | 94 | 291 | 91 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 35 | NA | -9999 | 164 | 388 | 164 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 22.6 | 374 | PASS | -9999 | 137 | 356 | 135 | PASS | 0 | 40 | QC1 | Probe | 110 | 109 | 109 | PASS | QC2 | Probe | 102 | 102 | 102 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.2 | 47.4 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 69.8 | 41.8 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.4 | 150.2 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 20.7 | -16.8 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_01_1_16+ | 614332205 | 92303 | Positive Control 1 | 839032 | 676886 | 6.8 | 2025-10-26T01:54:19+0200 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 24.1 | 162 | NA | -9999 | 197 | 428 | 191 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 26.2 | 120 | NA | -9999 | 179 | 421 | 176 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 23.4 | 427 | PASS | -9999 | 237 | 585 | 234 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24.6 | 185 | PASS | -9999 | 90 | 273 | 87 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 34 | NA | -9999 | 160 | 379 | 160 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 22.4 | 393 | PASS | -9999 | 131 | 341 | 128 | PASS | 0 | 40 | QC1 | Probe | 104 | 104 | 104 | PASS | QC2 | Probe | 107 | 107 | 107 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.4 | 51.4 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70.1 | 46.1 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.6 | 140.5 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 15.5 | -9.3 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_01_1_17+ | 614332207 | 92303 | Positive Control 1 | 839032 | 210024182 | 6.8 | 2025-10-26T01:54:46+0200 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 25 | 141 | NA | -9999 | 168 | 359 | 165 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 26.8 | 123 | NA | -9999 | 157 | 372 | 156 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 24.6 | 416 | PASS | -9999 | 188 | 462 | 187 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 26.1 | 166 | PASS | -9999 | 101 | 304 | 99 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 35 | NA | -9999 | 153 | 357 | 152 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 23.2 | 386 | PASS | -9999 | 126 | 318 | 123 | PASS | 0 | 40 | QC1 | Probe | 108 | 107 | 107 | PASS | QC2 | Probe | 89 | 89 | 89 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.3 | 41.5 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 69.9 | 42.8 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.5 | 138.9 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 20.3 | -7.3 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_01_1_18+ | 614332209 | 92303 | Positive Control 1 | 839032 | 210077690 | 6.8 | 2025-10-26T01:55:09+0200 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 23.2 | 143 | NA | -9999 | 151 | 304 | 143 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 25.9 | 119 | NA | -9999 | 174 | 405 | 169 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 23 | 429 | PASS | -9999 | 218 | 526 | 209 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24 | 193 | PASS | -9999 | 95 | 293 | 92 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 29 | NA | -9999 | 158 | 356 | 153 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 21.6 | 439 | PASS | -9999 | 121 | 321 | 116 | PASS | 0 | 40 | QC1 | Probe | 102 | 101 | 101 | PASS | QC2 | Probe | 107 | 106 | 106 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.3 | 41.4 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 69.9 | 45 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.5 | 144.8 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 22 | -11.9 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_02_1_11X | 614334666 | 92303 | Positive Control 1 | 839032 | 681294 | 6.8 | 26/10/25 03:31:32 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 24.2 | 160 | NA | -9999 | 209 | 464 | 204 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 25.8 | 125 | NA | -9999 | 215 | 512 | 213 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 23.8 | 410 | PASS | -9999 | 296 | 731 | 293 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 23.8 | 205 | PASS | -9999 | 142 | 431 | 139 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 33 | NA | -9999 | 171 | 423 | 170 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 22.1 | 426 | PASS | -9999 | 131 | 347 | 129 | PASS | 0 | 40 | QC1 | Probe | 115 | 115 | 114 | PASS | QC2 | Probe | 112 | 112 | 112 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.3 | 53.4 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70.1 | 65.7 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.6 | 167.1 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 18.3 | -10.1 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_02_1_12X | 614334122 | 92303 | Positive Control 1 | 839032 | 210058622 | 6.8 | 26/10/25 03:31:57 | Aborted | ERROR | SPC | SPC | 1,2,3,4 | NA | 23.9 | 161 | NA | -9999 | 151 | 330 | 147 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 25.7 | 110 | NA | -9999 | 185 | 443 | 184 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 22.9 | 442 | PASS | -9999 | 243 | 595 | 240 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24.1 | 181 | PASS | -9999 | 115 | 351 | 113 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 37 | NA | -9999 | 161 | 379 | 160 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 21.7 | 417 | PASS | -9999 | 184 | 477 | 181 | PASS | 0 | 40 | QC1 | Probe | 118 | 117 | 117 | PASS | QC2 | Probe | 101 | 101 | 101 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.3 | 50.9 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 69.9 | 44.3 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.5 | 191.6 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 17.4 | -9.8 | Error 5007: [CT1] probe check failed. Probe check value of 17 for reading number 2 was below the minimum of 110¦Error 5007: [SPC] probe check failed. Probe check value of 19 for reading number 2 was below the minimum of 119 |  |  |  |  | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_02_1_13X | 614334946 | 92303 | Positive Control 1 | 839032 | 652304 | 6.8 | 26/10/25 03:32:23 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 24.5 | 145 | NA | -9999 | 221 | 484 | 216 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 25.5 | 118 | NA | -9999 | 203 | 487 | 200 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 23.6 | 398 | PASS | -9999 | 264 | 656 | 260 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24.5 | 178 | PASS | -9999 | 119 | 376 | 117 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 36 | NA | -9999 | 185 | 444 | 185 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 22.1 | 402 | PASS | -9999 | 132 | 348 | 130 | PASS | 0 | 40 | QC1 | Probe | 117 | 117 | 117 | PASS | QC2 | Probe | 117 | 117 | 116 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.4 | 54.6 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70.1 | 50.7 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.6 | 149.6 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 19.7 | -9.6 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_02_1_14X | 614335244 | 92303 | Positive Control 1 | 839032 | 653259 | 6.8 | 26/10/25 03:32:47 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 24.2 | 162 | NA | -9999 | 234 | 503 | 225 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 24.9 | 130 | NA | -9999 | 204 | 484 | 200 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 23 | 444 | PASS | -9999 | 244 | 603 | 240 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24.2 | 182 | PASS | -9999 | 126 | 375 | 124 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 39 | NA | -9999 | 187 | 444 | 186 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 21.7 | 419 | PASS | -9999 | 150 | 391 | 148 | PASS | 0 | 40 | QC1 | Probe | 123 | 122 | 122 | PASS | QC2 | Probe | 113 | 113 | 113 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.3 | 50.9 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70.1 | 56.2 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.6 | 179.2 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 19.2 | -9 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_02_1_15+ | 614333907 | 92303 | Positive Control 1 | 839032 | 627453 | 6.8 | 26/10/25 03:33:13 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 25.2 | 147 | NA | -9999 | 248 | 516 | 243 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 26.7 | 117 | NA | -9999 | 198 | 468 | 196 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 23.8 | 421 | PASS | -9999 | 265 | 646 | 262 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24.9 | 179 | PASS | -9999 | 98 | 299 | 96 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 32 | NA | -9999 | 169 | 399 | 168 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 22.4 | 415 | PASS | -9999 | 129 | 334 | 128 | PASS | 0 | 40 | QC1 | Probe | 102 | 102 | 102 | PASS | QC2 | Probe | 102 | 101 | 101 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.1 | 55.1 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 69.8 | 42.5 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.3 | 146.8 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 15.8 | -9 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_02_1_16+ | 614334807 | 92303 | Positive Control 1 | 839032 | 782120 | 6.8 | 26/10/25 03:33:38 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 24.2 | 150 | NA | -9999 | 205 | 410 | 199 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 25.4 | 116 | NA | -9999 | 199 | 468 | 197 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 23 | 407 | PASS | -9999 | 280 | 680 | 278 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24.1 | 181 | PASS | -9999 | 103 | 319 | 101 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 37 | NA | -9999 | 168 | 396 | 168 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 21.8 | 406 | PASS | -9999 | 150 | 398 | 147 | PASS | 0 | 40 | QC1 | Probe | 114 | 114 | 114 | PASS | QC2 | Probe | 119 | 119 | 119 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.2 | 61.1 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 69.8 | 46.6 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.4 | 171.8 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 24.8 | -15.1 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_02_1_17+ | 614335071 | 92303 | Positive Control 1 | 839032 | 676886 | 6.8 | 26/10/25 03:34:24 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 24 | 158 | NA | -9999 | 201 | 439 | 196 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 26.2 | 124 | NA | -9999 | 186 | 439 | 184 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 23.4 | 434 | PASS | -9999 | 239 | 591 | 237 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24.4 | 190 | PASS | -9999 | 93 | 280 | 91 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 38 | NA | -9999 | 156 | 376 | 156 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 22.1 | 418 | PASS | -9999 | 128 | 334 | 126 | PASS | 0 | 40 | QC1 | Probe | 107 | 106 | 106 | PASS | QC2 | Probe | 100 | 100 | 100 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.3 | 49.3 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70 | 47 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.5 | 156.8 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 17.1 | -9.8 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_02_1_18+ | 614334181 | 92303 | Positive Control 1 | 839032 | 210034622 | 6.8 | 26/10/25 03:34:03 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 24.4 | 153 | NA | -9999 | 155 | 347 | 149 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 26.4 | 118 | NA | -9999 | 165 | 396 | 164 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 24.3 | 376 | PASS | -9999 | 222 | 547 | 221 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24.9 | 180 | PASS | -9999 | 91 | 284 | 89 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 30 | NA | -9999 | 154 | 358 | 154 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 22.7 | 370 | PASS | -9999 | 123 | 321 | 121 | PASS | 0 | 40 | QC1 | Probe | 114 | 114 | 114 | PASS | QC2 | Probe | 99 | 99 | 99 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.2 | 49.5 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 69.8 | 49 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.3 | 144.2 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 19.6 | -9.7 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_03_1_11X | 614338636 | 92303 | Positive Control 1 | 839032 | 210016126 | 6.8 | 2025-10-26T04:53:46+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 24 | 147 | NA | -9999 | 234 | 495 | 231 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 26 | 125 | NA | -9999 | 192 | 459 | 191 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 23.5 | 437 | PASS | -9999 | 280 | 697 | 280 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24.4 | 190 | PASS | -9999 | 111 | 334 | 109 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 36 | NA | -9999 | 171 | 408 | 171 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 21.8 | 436 | PASS | -9999 | 145 | 377 | 143 | PASS | 0 | 40 | QC1 | Probe | 118 | 117 | 117 | PASS | QC2 | Probe | 111 | 111 | 110 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.4 | 58.8 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70.1 | 51 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.6 | 172 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 16.7 | -9.3 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_03_1_12X | 614337720 | 92303 | Positive Control 1 | 839032 | 736872 | 6.8 | 2025-10-26T04:55:01+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 25.2 | 177 | NA | -9999 | 195 | 443 | 188 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 26.8 | 129 | NA | -9999 | 233 | 548 | 229 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 24.6 | 441 | PASS | -9999 | 285 | 708 | 282 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 25.6 | 190 | PASS | -9999 | 114 | 353 | 112 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 33 | NA | -9999 | 190 | 464 | 189 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 23.5 | 427 | PASS | -9999 | 147 | 393 | 144 | PASS | 0 | 40 | QC1 | Probe | 112 | 112 | 111 | PASS | QC2 | Probe | 115 | 115 | 115 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.4 | 48.7 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70.2 | 43.8 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.7 | 180.1 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 21.1 | -13.7 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_03_1_13X | 614338296 | 92303 | Positive Control 1 | 839032 | 739871 | 6.8 | 2025-10-26T04:56:13+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 25.4 | 162 | NA | -9999 | 197 | 430 | 190 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 26.2 | 126 | NA | -9999 | 191 | 457 | 188 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 24 | 415 | PASS | -9999 | 261 | 645 | 259 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 25.2 | 175 | PASS | -9999 | 109 | 332 | 107 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 30 | NA | -9999 | 166 | 395 | 165 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 22.8 | 385 | PASS | -9999 | 140 | 376 | 138 | PASS | 0 | 40 | QC1 | Probe | 82 | 81 | 81 | PASS | QC2 | Probe | 100 | 100 | 100 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.5 | 53 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70 | 44.1 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.5 | 163 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 19.3 | -6.5 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_03_1_14AX | 614338044 | 92303 | Positive Control 1 | 839032 | 656159 | 6.8 | 2025-10-26T08:58:18+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 24.4 | 154 | NA | -9999 | 181 | 393 | 177 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 25.9 | 121 | NA | -9999 | 207 | 494 | 204 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 23.7 | 410 | PASS | -9999 | 262 | 648 | 258 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24.8 | 177 | PASS | -9999 | 114 | 344 | 111 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 31 | NA | -9999 | 163 | 391 | 163 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 22.4 | 411 | PASS | -9999 | 131 | 341 | 128 | PASS | 0 | 40 | QC1 | Probe | 106 | 105 | 105 | PASS | QC2 | Probe | 108 | 108 | 108 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.4 | 55.4 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70 | 48.9 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.5 | 152 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 20 | -5.6 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_03_1_15A+ | 614338046 | 92303 | Positive Control 1 | 839032 | 210058622 | 6.8 | 2025-10-26T08:59:18+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 24 | 155 | NA | -9999 | 146 | 344 | 142 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 27.5 | 125 | NA | -9999 | 202 | 478 | 199 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 25.4 | 393 | PASS | -9999 | 253 | 634 | 250 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 26 | 182 | PASS | -9999 | 101 | 309 | 99 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 32 | NA | -9999 | 182 | 437 | 181 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 24.1 | 392 | PASS | -9999 | 131 | 347 | 129 | PASS | 0 | 40 | QC1 | Probe | 115 | 114 | 114 | PASS | QC2 | Probe | 122 | 122 | 122 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.4 | 41.2 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70.1 | 36.4 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.6 | 139.7 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 24.6 | -16 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_03_1_16A+ | 614338467 | 92303 | Positive Control 1 | 839032 | 652304 | 6.8 | 2025-10-26T09:00:18+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 23.4 | 160 | NA | -9999 | 192 | 422 | 186 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 26 | 130 | NA | -9999 | 207 | 491 | 203 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 23.4 | 435 | PASS | -9999 | 241 | 597 | 237 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24.4 | 188 | PASS | -9999 | 111 | 332 | 109 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 37 | NA | -9999 | 186 | 447 | 184 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 22.2 | 430 | PASS | -9999 | 133 | 343 | 129 | PASS | 0 | 40 | QC1 | Probe | 114 | 113 | 113 | PASS | QC2 | Probe | 129 | 129 | 129 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.4 | 44.7 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70.1 | 39.1 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.6 | 138.3 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 15.3 | -7.9 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_03_1_17+ | 614337607 | 92303 | Positive Control 1 | 839032 | 210003146 | 6.8 | 2025-10-26T04:58:16+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 24.3 | 162 | NA | -9999 | 185 | 398 | 179 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 24.7 | 137 | NA | -9999 | 173 | 411 | 170 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 22.5 | 449 | PASS | -9999 | 255 | 630 | 252 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 23.2 | 199 | PASS | -9999 | 101 | 309 | 99 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 28 | NA | -9999 | 157 | 376 | 157 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 20.7 | 452 | PASS | -9999 | 119 | 314 | 116 | PASS | 0 | 40 | QC1 | Probe | 98 | 97 | 97 | PASS | QC2 | Probe | 101 | 101 | 101 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.2 | 53.3 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 69.9 | 49.6 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.4 | 151.8 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 19.5 | -8.6 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_03_1_18+ | 614338203 | 92303 | Positive Control 1 | 839032 | 694834 | 6.8 | 2025-10-26T04:58:39+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 24.5 | 144 | NA | -9999 | 218 | 476 | 211 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 26.6 | 116 | NA | -9999 | 215 | 507 | 213 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 24.1 | 399 | PASS | -9999 | 284 | 702 | 282 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24.8 | 185 | PASS | -9999 | 124 | 385 | 122 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 31 | NA | -9999 | 181 | 436 | 181 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 22.5 | 412 | PASS | -9999 | 152 | 410 | 149 | PASS | 0 | 40 | QC1 | Probe | 124 | 123 | 123 | PASS | QC2 | Probe | 117 | 117 | 117 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.3 | 53 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70 | 51.3 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.5 | 169.1 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 16.8 | -10.4 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_04_1_11X | 614341522 | 92303 | Positive Control 1 | 839032 | 210016126 | 6.8 | 2025-10-26T08:52:27+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 23.4 | 164 | NA | -9999 | 234 | 496 | 232 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 25 | 131 | NA | -9999 | 198 | 471 | 197 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 22.5 | 459 | PASS | -9999 | 285 | 710 | 286 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 23.4 | 202 | PASS | -9999 | 113 | 337 | 111 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 38 | NA | -9999 | 176 | 417 | 177 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 21 | 450 | PASS | -9999 | 139 | 361 | 137 | PASS | 0 | 40 | QC1 | Probe | 121 | 120 | 120 | PASS | QC2 | Probe | 120 | 120 | 120 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.4 | 57.3 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70.1 | 49.4 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.6 | 151.2 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 15.3 | -7.6 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_04_1_12X | 614341214 | 92303 | Positive Control 1 | 839032 | 736872 | 6.8 | 2025-10-26T08:52:53+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 24.2 | 169 | NA | -9999 | 180 | 397 | 174 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 25.8 | 120 | NA | -9999 | 203 | 483 | 200 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 23.6 | 413 | PASS | -9999 | 255 | 633 | 255 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24.3 | 193 | PASS | -9999 | 104 | 327 | 102 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 33 | NA | -9999 | 182 | 421 | 181 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 22.1 | 408 | PASS | -9999 | 133 | 355 | 131 | PASS | 0 | 40 | QC1 | Probe | 119 | 118 | 118 | PASS | QC2 | Probe | 125 | 125 | 125 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.4 | 51.7 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70 | 47.1 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.5 | 141.2 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 21.1 | -12.1 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_04_1_13X | 614340896 | 92303 | Positive Control 1 | 839032 | 739871 | 6.8 | 2025-10-26T08:53:17+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 23.9 | 155 | NA | -9999 | 194 | 422 | 190 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 25.5 | 124 | NA | -9999 | 183 | 436 | 181 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 23.2 | 425 | PASS | -9999 | 250 | 618 | 248 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24.1 | 187 | PASS | -9999 | 104 | 323 | 103 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 31 | NA | -9999 | 159 | 380 | 159 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 21.7 | 417 | PASS | -9999 | 137 | 370 | 135 | PASS | 0 | 40 | QC1 | Probe | 116 | 116 | 116 | PASS | QC2 | Probe | 108 | 108 | 108 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.4 | 61.1 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70 | 57.6 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.5 | 169.2 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 18.4 | -8.7 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_04_1_14AX | 614341834 | 92303 | Positive Control 1 | 839032 | 210018949 | 6.8 | 2025-10-26T13:13:00+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 24.8 | 158 | NA | -9999 | 205 | 441 | 199 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 26.9 | 110 | NA | -9999 | 215 | 509 | 211 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 24.5 | 397 | PASS | -9999 | 282 | 688 | 278 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 25.2 | 183 | PASS | -9999 | 109 | 336 | 106 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 31 | NA | -9999 | 160 | 353 | 159 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 22.9 | 402 | PASS | -9999 | 147 | 384 | 144 | PASS | 0 | 40 | QC1 | Probe | 110 | 110 | 109 | PASS | QC2 | Probe | 110 | 110 | 110 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.4 | 51.7 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70 | 40.9 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.5 | 152.4 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 20.1 | -6.5 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_04_1_15+ | 614341037 | 92303 | Positive Control 1 | 839032 | 210059038 | 6.8 | 2025-10-26T08:53:42+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 26.2 | 139 | NA | -9999 | 145 | 323 | 141 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 26.9 | 100 | NA | -9999 | 186 | 441 | 183 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 25.1 | 361 | PASS | -9999 | 231 | 574 | 228 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 25.7 | 169 | PASS | -9999 | 92 | 284 | 89 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 26 | NA | -9999 | 172 | 397 | 170 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 23.6 | 354 | PASS | -9999 | 137 | 360 | 134 | PASS | 0 | 40 | QC1 | Probe | 125 | 124 | 123 | PASS | QC2 | Probe | 103 | 103 | 103 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.4 | 48.7 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 69.9 | 34.8 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.4 | 127.7 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 19.8 | -10.3 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_04_1_16+ | 614341409 | 92303 | Positive Control 1 | 839032 | 210012379 | 6.8 | 2025-10-26T08:54:10+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 25.7 | 148 | NA | -9999 | 186 | 409 | 180 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 26.4 | 119 | NA | -9999 | 205 | 488 | 203 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 24.7 | 397 | PASS | -9999 | 243 | 604 | 240 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 25.7 | 171 | PASS | -9999 | 112 | 337 | 110 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 32 | NA | -9999 | 185 | 448 | 184 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 23.2 | 381 | PASS | -9999 | 136 | 352 | 133 | PASS | 0 | 40 | QC1 | Probe | 109 | 109 | 108 | PASS | QC2 | Probe | 116 | 115 | 115 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.5 | 50.8 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70.1 | 52.2 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.6 | 145.6 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 18.2 | -9 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_04_1_17+ | 614340747 | 92303 | Positive Control 1 | 839032 | 210003146 | 6.8 | 2025-10-26T08:54:32+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 23.5 | 163 | NA | -9999 | 174 | 372 | 168 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 24.8 | 134 | NA | -9999 | 166 | 397 | 164 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 23.1 | 441 | PASS | -9999 | 240 | 591 | 239 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 23.7 | 203 | PASS | -9999 | 95 | 305 | 94 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 29 | NA | -9999 | 154 | 365 | 155 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 21.4 | 429 | PASS | -9999 | 121 | 319 | 119 | PASS | 0 | 40 | QC1 | Probe | 102 | 102 | 101 | PASS | QC2 | Probe | 106 | 106 | 106 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.3 | 54.9 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 69.9 | 49 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.4 | 129 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 21.1 | -10.2 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_04_1_18+ | 614341905 | 92303 | Positive Control 1 | 839032 | 694834 | 6.8 | 2025-10-26T08:54:58+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 24.8 | 140 | NA | -9999 | 207 | 443 | 201 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 26.1 | 104 | NA | -9999 | 198 | 463 | 194 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 23.9 | 373 | PASS | -9999 | 255 | 623 | 253 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24.5 | 177 | PASS | -9999 | 119 | 362 | 117 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 29 | NA | -9999 | 169 | 390 | 168 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 22.2 | 374 | PASS | -9999 | 148 | 395 | 145 | PASS | 0 | 40 | QC1 | Probe | 131 | 130 | 130 | PASS | QC2 | Probe | 120 | 119 | 119 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.2 | 51.5 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 69.9 | 52 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.4 | 162.3 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 18.1 | -8.3 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_05_1_11X | 614344758 | 92303 | Positive Control 1 | 839032 | 681294 | 6.8 | 2025-10-26T08:55:24+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 24.3 | 154 | NA | -9999 | 212 | 458 | 206 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 26.1 | 116 | NA | -9999 | 212 | 505 | 209 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 24.1 | 400 | PASS | -9999 | 289 | 709 | 285 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24.2 | 198 | PASS | -9999 | 142 | 433 | 138 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 32 | NA | -9999 | 167 | 410 | 167 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 22.3 | 412 | PASS | -9999 | 142 | 375 | 139 | PASS | 0 | 40 | QC1 | Probe | 123 | 122 | 122 | PASS | QC2 | Probe | 120 | 120 | 120 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.4 | 54.6 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70 | 53.9 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.6 | 148.8 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 18 | -9 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_05_1_12X | 614345044 | 92303 | Positive Control 1 | 839032 | 210058622 | 6.8 | 2025-10-26T08:55:49+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 25.3 | 170 | NA | -9999 | 166 | 360 | 161 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 26 | 125 | NA | -9999 | 202 | 481 | 201 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 23.6 | 449 | PASS | -9999 | 251 | 618 | 249 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24.8 | 188 | PASS | -9999 | 118 | 361 | 117 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 33 | NA | -9999 | 164 | 393 | 164 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 22.5 | 418 | PASS | -9999 | 154 | 406 | 153 | PASS | 0 | 40 | QC1 | Probe | 117 | 116 | 116 | PASS | QC2 | Probe | 96 | 95 | 95 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.3 | 45.6 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70.1 | 45.5 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.5 | 177.5 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 18.7 | -11.6 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_05_1_13X | 614345292 | 92303 | Positive Control 1 | 839032 | 656159 | 6.8 | 2025-10-26T08:56:11+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 23.4 | 150 | NA | -9999 | 189 | 409 | 184 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 26.1 | 127 | NA | -9999 | 211 | 501 | 207 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 23.5 | 439 | PASS | -9999 | 267 | 663 | 263 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24.5 | 191 | PASS | -9999 | 113 | 340 | 110 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 29 | NA | -9999 | 170 | 399 | 169 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 22.4 | 425 | PASS | -9999 | 134 | 350 | 131 | PASS | 0 | 40 | QC1 | Probe | 108 | 108 | 107 | PASS | QC2 | Probe | 119 | 119 | 119 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.3 | 53.3 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70 | 48.4 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.5 | 158.4 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 16.6 | -8.4 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_05_1_14X | 614344398 | 92303 | Positive Control 1 | 839032 | 652304 | 6.8 | 2025-10-26T08:56:39+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 23.6 | 155 | NA | -9999 | 221 | 484 | 214 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 25.4 | 136 | NA | -9999 | 203 | 484 | 200 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 23.1 | 423 | PASS | -9999 | 263 | 652 | 259 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24 | 203 | PASS | -9999 | 119 | 372 | 117 | PASS | 0 | 40 | rpoB3 | Target | 3 | POS | 40 | 40 | PASS | -9999 | 188 | 448 | 188 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 21.7 | 512 | PASS | -9999 | 129 | 338 | 127 | PASS | 0 | 40 | QC1 | Probe | 115 | 115 | 114 | PASS | QC2 | Probe | 123 | 122 | 122 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.5 | 49.2 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 69.9 | 38 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.5 | 111.3 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 18 | -10.8 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_05_1_15+ | 614344525 | 92303 | Positive Control 1 | 839032 | 653259 | 6.8 | 2025-10-26T08:57:03+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 24 | 165 | NA | -9999 | 237 | 509 | 229 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 25 | 130 | NA | -9999 | 209 | 493 | 205 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 22.5 | 455 | PASS | -9999 | 252 | 621 | 248 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 23.9 | 185 | PASS | -9999 | 125 | 374 | 123 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 39 | NA | -9999 | 187 | 450 | 187 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 21.3 | 433 | PASS | -9999 | 158 | 412 | 156 | PASS | 0 | 40 | QC1 | Probe | 129 | 129 | 128 | PASS | QC2 | Probe | 123 | 123 | 123 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.3 | 51.5 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70 | 53.9 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.6 | 183.9 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 17.6 | -7 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_05_1_16+ | 614344867 | 92303 | Positive Control 1 | 839032 | 627453 | 6.8 | 2025-10-26T08:57:29+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 24.1 | 142 | NA | -9999 | 249 | 520 | 247 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 26.6 | 116 | NA | -9999 | 200 | 473 | 199 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 23.6 | 426 | PASS | -9999 | 261 | 637 | 258 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24.6 | 185 | PASS | -9999 | 99 | 298 | 98 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 35 | NA | -9999 | 172 | 395 | 171 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 22.3 | 416 | PASS | -9999 | 126 | 328 | 125 | PASS | 0 | 40 | QC1 | Probe | 107 | 106 | 107 | PASS | QC2 | Probe | 110 | 109 | 109 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.3 | 52.2 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 69.9 | 37.5 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.4 | 136.9 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 18.1 | -8.1 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_05_1_17+ | 614344227 | 92303 | Positive Control 1 | 839032 | 676886 | 6.8 | 2025-10-26T08:57:53+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 25.2 | 148 | NA | -9999 | 206 | 448 | 202 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 25.9 | 115 | NA | -9999 | 186 | 437 | 184 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 23.2 | 420 | PASS | -9999 | 238 | 590 | 237 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24.2 | 186 | PASS | -9999 | 92 | 278 | 91 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 34 | NA | -9999 | 157 | 375 | 157 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 22 | 409 | PASS | -9999 | 132 | 343 | 130 | PASS | 0 | 40 | QC1 | Probe | 103 | 102 | 102 | PASS | QC2 | Probe | 99 | 99 | 99 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.4 | 51.1 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70 | 42.4 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.6 | 152.5 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 16.5 | -8.6 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_05_1_18+ | 614343985 | 92303 | Positive Control 1 | 839032 | 210024182 | 6.8 | 2025-10-26T08:58:18+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 24.7 | 160 | NA | -9999 | 177 | 383 | 174 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 26.3 | 131 | NA | -9999 | 178 | 420 | 175 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 23.5 | 462 | PASS | -9999 | 206 | 508 | 202 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 25 | 183 | PASS | -9999 | 104 | 313 | 102 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 40 | NA | -9999 | 157 | 375 | 157 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 22.4 | 429 | PASS | -9999 | 133 | 341 | 131 | PASS | 0 | 40 | QC1 | Probe | 108 | 107 | 107 | PASS | QC2 | Probe | 93 | 93 | 93 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.4 | 37.7 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70 | 39.3 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.6 | 145.5 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 20 | -9.5 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_06_1_11X | 614348326 | 92303 | Positive Control 1 | 839032 | 210019897 | 6.8 | 2025-10-26T13:13:26+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 24.5 | 152 | NA | -9999 | 188 | 407 | 179 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 25.6 | 124 | NA | -9999 | 171 | 405 | 166 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 23.4 | 431 | PASS | -9999 | 224 | 551 | 220 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24.4 | 178 | PASS | -9999 | 108 | 331 | 105 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 33 | NA | -9999 | 176 | 417 | 175 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 21.7 | 400 | PASS | -9999 | 143 | 368 | 139 | PASS | 0 | 40 | QC1 | Probe | 112 | 112 | 111 | PASS | QC2 | Probe | 110 | 110 | 109 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.2 | 57.2 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70 | 64.1 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.5 | 178.9 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 30.3 | -18.1 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_06_1_12X | 614348666 | 92303 | Positive Control 1 | 839032 | 730217 | 6.8 | 2025-10-26T13:13:51+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 24.1 | 170 | NA | -9999 | 241 | 512 | 231 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 25.9 | 131 | NA | -9999 | 209 | 490 | 205 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 23.1 | 444 | PASS | -9999 | 278 | 676 | 274 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24.2 | 199 | PASS | -9999 | 98 | 328 | 96 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 27 | NA | -9999 | 152 | 356 | 152 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 21.9 | 427 | PASS | -9999 | 142 | 376 | 139 | PASS | 0 | 40 | QC1 | Probe | 116 | 115 | 115 | PASS | QC2 | Probe | 110 | 110 | 110 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.3 | 50.8 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70 | 45.8 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.5 | 172 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 25.7 | -16.4 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_06_1_13AAX | 614345588 | 92303 | Positive Control 1 | 839032 | 690802 | 6.8 | 2025-10-26T17:29:12+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 23.5 | 161 | NA | -9999 | 240 | 499 | 234 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 25.7 | 117 | NA | -9999 | 232 | 540 | 229 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 22.3 | 459 | PASS | -9999 | 299 | 736 | 297 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 23.5 | 194 | PASS | -9999 | 123 | 369 | 121 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 37 | NA | -9999 | 189 | 428 | 188 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 21.2 | 431 | PASS | -9999 | 160 | 419 | 157 | PASS | 0 | 40 | QC1 | Probe | 136 | 135 | 135 | PASS | QC2 | Probe | 129 | 128 | 128 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.2 | 61.6 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 69.9 | 57.2 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.5 | 190 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 17.9 | -7.3 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_06_1_14X | 614346260 | 92303 | Positive Control 1 | 839032 | 210026044 | 6.8 | 2025-10-26T13:14:40+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 24.5 | 166 | NA | -9999 | 225 | 495 | 216 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 26.5 | 130 | NA | -9999 | 176 | 423 | 173 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 24 | 435 | PASS | -9999 | 270 | 670 | 267 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 25.1 | 182 | PASS | -9999 | 89 | 273 | 87 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 38 | NA | -9999 | 183 | 443 | 182 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 22.5 | 452 | PASS | -9999 | 94 | 250 | 92 | PASS | 0 | 40 | QC1 | Probe | 78 | 77 | 77 | PASS | QC2 | Probe | 112 | 112 | 112 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.4 | 46.3 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70.1 | 35.2 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.5 | 115.3 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 19 | -6.8 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_06_1_15+ | 614345727 | 92303 | Positive Control 1 | 839032 | 638845 | 6.8 | 2025-10-26T13:15:05+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 25.1 | 157 | NA | -9999 | 225 | 488 | 217 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 26.1 | 127 | NA | -9999 | 204 | 487 | 202 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 23.8 | 427 | PASS | -9999 | 278 | 694 | 275 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24.8 | 187 | PASS | -9999 | 108 | 329 | 106 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 36 | NA | -9999 | 180 | 440 | 180 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 22.4 | 429 | PASS | -9999 | 149 | 389 | 146 | PASS | 0 | 40 | QC1 | Probe | 111 | 111 | 110 | PASS | QC2 | Probe | 108 | 108 | 107 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.4 | 53.9 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70 | 46.9 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.5 | 181.4 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 19.4 | -7.9 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_06_1_16A+_D06 | 614348455 | 92303 | Positive Control 1 | 839032 | 210035802 | 6.8 | 2025-10-26T13:15:30+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 24.4 | 166 | NA | -9999 | 229 | 484 | 221 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 25.5 | 126 | NA | -9999 | 190 | 449 | 188 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 23.1 | 444 | PASS | -9999 | 241 | 588 | 238 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24.2 | 190 | PASS | -9999 | 106 | 333 | 104 | PASS | 0 | 40 | rpoB3 | Target | 3 | POS | 35 | 47 | PASS | -9999 | 168 | 407 | 168 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 21.9 | 435 | PASS | -9999 | 137 | 358 | 135 | PASS | 0 | 40 | QC1 | Probe | 110 | 110 | 110 | PASS | QC2 | Probe | 109 | 108 | 108 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.3 | 49.7 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70 | 49.5 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.5 | 158.9 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 28.4 | -19 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_06_1_17+ | 614347939 | 92303 | Positive Control 1 | 839032 | 880236 | 6.8 | 2025-10-26T13:15:56+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 23.7 | 162 | NA | -9999 | 188 | 411 | 180 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 26.1 | 127 | NA | -9999 | 190 | 448 | 186 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 23.9 | 424 | PASS | -9999 | 239 | 593 | 235 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24.9 | 184 | PASS | -9999 | 109 | 331 | 107 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 30 | NA | -9999 | 160 | 382 | 159 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 22.7 | 396 | PASS | -9999 | 139 | 361 | 136 | PASS | 0 | 40 | QC1 | Probe | 116 | 116 | 115 | PASS | QC2 | Probe | 109 | 108 | 108 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.4 | 41 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70 | 33.6 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.6 | 137.1 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 26.7 | -17.1 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_06_1_18+ | 614346409 | 92303 | Positive Control 1 | 839032 | 210016126 | 6.8 | 2025-10-26T13:16:18+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 24.5 | 159 | NA | -9999 | 240 | 513 | 237 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 25.8 | 131 | NA | -9999 | 205 | 485 | 203 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 23.3 | 462 | PASS | -9999 | 291 | 727 | 291 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24.3 | 198 | PASS | -9999 | 113 | 341 | 112 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 36 | NA | -9999 | 181 | 429 | 180 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 21.9 | 451 | PASS | -9999 | 148 | 384 | 146 | PASS | 0 | 40 | QC1 | Probe | 119 | 119 | 119 | PASS | QC2 | Probe | 111 | 111 | 111 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.5 | 59.4 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70.1 | 48.8 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.6 | 176.1 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 17.4 | -7.1 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_07_1_11X | 614350058 | 92303 | Positive Control 1 | 839032 | 210018949 | 6.8 | 2025-10-26T15:00:13+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 24.6 | 168 | NA | -9999 | 183 | 394 | 178 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 25.3 | 121 | NA | -9999 | 196 | 467 | 194 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 22.6 | 432 | PASS | -9999 | 263 | 646 | 259 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 23.5 | 196 | PASS | -9999 | 105 | 323 | 103 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 39 | NA | -9999 | 146 | 331 | 145 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 21.2 | 423 | PASS | -9999 | 140 | 362 | 137 | PASS | 0 | 40 | QC1 | Probe | 113 | 113 | 112 | PASS | QC2 | Probe | 97 | 96 | 96 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.4 | 59.3 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70 | 53.8 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.6 | 163.9 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 26.2 | -13.5 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_07_1_12X | 614350324 | 92303 | Positive Control 1 | 839032 | 210019897 | 6.8 | 2025-10-26T15:00:39+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 24.2 | 155 | NA | -9999 | 208 | 461 | 201 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 27.1 | 125 | NA | -9999 | 193 | 456 | 189 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 24.2 | 433 | PASS | -9999 | 244 | 608 | 240 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 25.2 | 180 | PASS | -9999 | 116 | 350 | 114 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 36 | NA | -9999 | 186 | 448 | 185 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 22.7 | 413 | PASS | -9999 | 152 | 395 | 148 | PASS | 0 | 40 | QC1 | Probe | 123 | 122 | 122 | PASS | QC2 | Probe | 116 | 116 | 116 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.4 | 45.2 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70.1 | 41.6 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.6 | 171.4 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 24.4 | -13.5 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_07_1_13X_D2B,11,05 | 614349478 | 92303 | Positive Control 1 | 839032 | 730217 | 6.8 | 2025-10-26T15:01:05+0100 | Aborted | ERROR | SPC | SPC | 1,2,3,4 | NA | 23.3 | 157 | NA | -9999 | 226 | 470 | 220 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 26.8 | 118 | NA | -9999 | 196 | 462 | 194 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 24.4 | 412 | PASS | -9999 | 257 | 629 | 256 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 25.3 | 190 | PASS | -9999 | 94 | 316 | 92 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 17 | NA | -9999 | 148 | 332 | 147 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 23.3 | 379 | PASS | -9999 | 142 | 375 | 140 | PASS | 0 | 40 | QC1 | Probe | 121 | 121 | 121 | PASS | QC2 | Probe | 107 | 106 | 106 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.3 | 48.6 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70 | 43.6 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.6 | 165.5 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 26 | -16.9 | Error 5006: [QC-2] probe check failed. Probe check value of 247 for reading number 1 was above the maximum of 223¦Error 5006: [QC-2] probe check failed. Probe check value of 247 for reading number 3 was above the maximum of 223 |  |  |  |  | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_07_1_14X | 614349732 | 92303 | Positive Control 1 | 839032 | 690802 | 6.8 | 2025-10-26T15:01:30+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 26.2 | 139 | NA | -9999 | 230 | 483 | 226 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 27.1 | 114 | NA | -9999 | 215 | 507 | 213 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 24.4 | 407 | PASS | -9999 | 285 | 699 | 284 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 25.9 | 167 | PASS | -9999 | 119 | 358 | 118 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 30 | NA | -9999 | 180 | 407 | 179 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 23.4 | 368 | PASS | -9999 | 145 | 383 | 143 | PASS | 0 | 40 | QC1 | Probe | 121 | 121 | 121 | PASS | QC2 | Probe | 115 | 115 | 115 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.1 | 53.4 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 69.8 | 43.2 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.4 | 159.3 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 24 | -13.1 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_07_1_15+ | 614349303 | 92303 | Positive Control 1 | 839032 | 210026044 | 6.8 | 2025-10-26T15:01:55+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 23.7 | 160 | NA | -9999 | 201 | 439 | 189 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 26.6 | 125 | NA | -9999 | 153 | 365 | 149 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 24.9 | 396 | PASS | -9999 | 241 | 594 | 236 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 25.7 | 175 | PASS | -9999 | 80 | 245 | 78 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 36 | NA | -9999 | 171 | 399 | 169 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 23 | 415 | PASS | -9999 | 84 | 221 | 81 | PASS | 0 | 40 | QC1 | Probe | 72 | 72 | 71 | PASS | QC2 | Probe | 110 | 110 | 110 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.3 | 45.7 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 69.9 | 31.8 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.4 | 107.1 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 24.7 | -13.5 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_07_1_16+ | 614350191 | 92303 | Positive Control 1 | 839032 | 638845 | 6.8 | 2025-10-26T15:02:20+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 24.4 | 159 | NA | -9999 | 213 | 459 | 206 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 26.9 | 122 | NA | -9999 | 200 | 478 | 198 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 24.8 | 414 | PASS | -9999 | 267 | 666 | 263 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 25.9 | 177 | PASS | -9999 | 106 | 324 | 105 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 32 | NA | -9999 | 177 | 426 | 177 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 23.6 | 407 | PASS | -9999 | 145 | 381 | 142 | PASS | 0 | 40 | QC1 | Probe | 120 | 120 | 120 | PASS | QC2 | Probe | 115 | 115 | 115 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.4 | 53.5 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 69.9 | 44.2 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.5 | 174.6 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 30.1 | -23.6 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_07_1_17+ | 614349857 | 92303 | Positive Control 1 | 839032 | 210035802 | 6.8 | 2025-10-26T15:02:46+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 24 | 170 | NA | -9999 | 232 | 496 | 223 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 26 | 131 | NA | -9999 | 185 | 435 | 181 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 23.3 | 448 | PASS | -9999 | 238 | 588 | 235 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24.4 | 192 | PASS | -9999 | 108 | 330 | 106 | PASS | 0 | 40 | rpoB3 | Target | 3 | POS | 35.2 | 49 | PASS | -9999 | 173 | 408 | 172 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 22.2 | 406 | PASS | -9999 | 137 | 356 | 134 | PASS | 0 | 40 | QC1 | Probe | 102 | 102 | 102 | PASS | QC2 | Probe | 106 | 106 | 106 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.4 | 50.5 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70.1 | 48.8 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.6 | 156.8 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 22.7 | -11.5 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_07_1_18+ | 614349093 | 92303 | Positive Control 1 | 839032 | 880236 | 6.8 | 2025-10-26T15:03:10+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 23.1 | 159 | NA | -9999 | 174 | 373 | 164 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 26.5 | 119 | NA | -9999 | 171 | 410 | 167 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 23.5 | 416 | PASS | -9999 | 216 | 536 | 213 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24.6 | 179 | PASS | -9999 | 101 | 311 | 99 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 30 | NA | -9999 | 154 | 355 | 153 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 22.4 | 364 | PASS | -9999 | 129 | 339 | 127 | PASS | 0 | 40 | QC1 | Probe | 115 | 114 | 113 | PASS | QC2 | Probe | 111 | 110 | 110 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.3 | 47.7 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 69.9 | 40.1 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.4 | 131.6 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 31.3 | -15.6 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_08_1_11X | 614384828 | 92303 | Positive Control 1 | 839032 | 210018949 | 6.8 | 2025-10-26T17:29:12+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 24.9 | 146 | NA | -9999 | 199 | 438 | 195 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 27.7 | 116 | NA | -9999 | 206 | 487 | 203 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 26 | 375 | PASS | -9999 | 284 | 701 | 282 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 26.9 | 168 | PASS | -9999 | 109 | 336 | 107 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 24 | NA | -9999 | 149 | 351 | 148 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 24.6 | 373 | PASS | -9999 | 134 | 351 | 132 | PASS | 0 | 40 | QC1 | Probe | 122 | 122 | 122 | PASS | QC2 | Probe | 106 | 106 | 106 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.4 | 55.3 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70.2 | 47.6 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.6 | 155.2 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 17.6 | -7.7 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_08_1_12X | 614384186 | 92303 | Positive Control 1 | 839032 | 210019897 | 6.8 | 2025-10-26T17:30:02+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 25.2 | 145 | NA | -9999 | 197 | 427 | 189 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 26.6 | 111 | NA | -9999 | 180 | 427 | 176 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 24.6 | 402 | PASS | -9999 | 233 | 572 | 228 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 25.8 | 166 | PASS | -9999 | 113 | 341 | 111 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 33 | NA | -9999 | 178 | 427 | 177 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 23 | 395 | PASS | -9999 | 149 | 385 | 145 | PASS | 0 | 40 | QC1 | Probe | 126 | 125 | 125 | PASS | QC2 | Probe | 114 | 114 | 114 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.3 | 53.5 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 69.9 | 49.2 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.4 | 162.7 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 18.8 | -8.8 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_08_1_13X | 614352594 | 92303 | Positive Control 1 | 839032 | 730217 | 6.8 | 2025-10-26T17:31:16+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 24.6 | 143 | NA | -9999 | 248 | 524 | 243 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 28 | 115 | NA | -9999 | 207 | 488 | 204 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 25.7 | 383 | PASS | -9999 | 278 | 681 | 275 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 26.7 | 173 | PASS | -9999 | 98 | 325 | 96 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 18 | NA | -9999 | 157 | 371 | 155 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 24.6 | 375 | PASS | -9999 | 145 | 382 | 142 | PASS | 0 | 40 | QC1 | Probe | 112 | 112 | 112 | PASS | QC2 | Probe | 100 | 99 | 99 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.5 | 48.2 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70 | 34.8 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.6 | 159.5 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 21.8 | -12.5 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_08_1_14X_D2B | 614382746 | 92303 | Positive Control 1 | 839032 | 690802 | 6.8 | 2025-10-26T17:32:06+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 24.7 | 155 | NA | -9999 | 236 | 499 | 231 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 26.4 | 123 | NA | -9999 | 224 | 526 | 221 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 23.5 | 439 | PASS | -9999 | 292 | 716 | 288 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24.7 | 185 | PASS | -9999 | 118 | 353 | 116 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 36 | NA | -9999 | 177 | 412 | 176 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 22.4 | 414 | PASS | -9999 | 147 | 387 | 144 | PASS | 0 | 40 | QC1 | Probe | 121 | 121 | 120 | PASS | QC2 | Probe | 120 | 119 | 119 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.2 | 59.3 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 69.9 | 51.4 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.4 | 181.1 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 26.6 | -10.8 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_08_1_15+ | 614383177 | 92303 | Positive Control 1 | 839032 | 210026044 | 6.8 | 2025-10-26T17:32:30+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 24.7 | 166 | NA | -9999 | 217 | 477 | 209 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 27.7 | 131 | NA | -9999 | 167 | 401 | 164 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 26 | 392 | PASS | -9999 | 259 | 644 | 256 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 27.4 | 158 | PASS | -9999 | 87 | 266 | 85 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 32 | NA | -9999 | 177 | 428 | 176 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 24.8 | 414 | PASS | -9999 | 95 | 251 | 93 | PASS | 0 | 40 | QC1 | Probe | 72 | 71 | 71 | PASS | QC2 | Probe | 99 | 99 | 99 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.4 | 44.8 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70.1 | 29.9 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.5 | 110.9 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 24.9 | -18.8 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_08_1_16+_D05 | 614384577 | 92303 | Positive Control 1 | 839032 | 638845 | 6.8 | 2025-10-26T17:32:55+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 25.1 | 143 | NA | -9999 | 221 | 460 | 213 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 25.9 | 111 | NA | -9999 | 199 | 475 | 197 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 24 | 393 | PASS | -9999 | 275 | 681 | 272 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24.8 | 172 | PASS | -9999 | 107 | 326 | 105 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 31 | NA | -9999 | 177 | 430 | 177 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 22.4 | 407 | PASS | -9999 | 149 | 393 | 147 | PASS | 0 | 40 | QC1 | Probe | 120 | 120 | 120 | PASS | QC2 | Probe | 117 | 117 | 116 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.2 | 64.2 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 69.8 | 51.3 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.3 | 185.8 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 19.7 | -7.1 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_08_1_17+ | 614352289 | 92303 | Positive Control 1 | 839032 | 210035802 | 6.8 | 2025-10-26T17:33:21+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 26.1 | 153 | NA | -9999 | 238 | 499 | 230 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 26.9 | 117 | NA | -9999 | 184 | 440 | 183 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 24.6 | 387 | PASS | -9999 | 240 | 591 | 239 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 25.6 | 173 | PASS | -9999 | 108 | 334 | 106 | PASS | 0 | 40 | rpoB3 | Target | 3 | POS | 37.8 | 43 | PASS | -9999 | 173 | 410 | 172 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 23.4 | 375 | PASS | -9999 | 137 | 358 | 135 | PASS | 0 | 40 | QC1 | Probe | 107 | 107 | 107 | PASS | QC2 | Probe | 102 | 101 | 101 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.3 | 49.8 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 69.9 | 48.5 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.5 | 153.7 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 33.9 | -15.8 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_08_1_18+ | 614383633 | 92303 | Positive Control 1 | 839032 | 880236 | 6.8 | 2025-10-26T17:33:44+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 24.3 | 154 | NA | -9999 | 175 | 381 | 166 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 25.8 | 125 | NA | -9999 | 176 | 417 | 173 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 24.2 | 403 | PASS | -9999 | 221 | 547 | 217 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 25.1 | 179 | PASS | -9999 | 108 | 325 | 105 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 27 | NA | -9999 | 149 | 358 | 148 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 23 | 370 | PASS | -9999 | 116 | 302 | 113 | PASS | 0 | 40 | QC1 | Probe | 116 | 116 | 115 | PASS | QC2 | Probe | 98 | 97 | 97 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.3 | 47.3 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70 | 53.7 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.5 | 138.5 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 31.9 | -21 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_09_1_11X | 614385278 | 92303 | Positive Control 1 | 839032 | 210018949 | 6.8 | 2025-10-26T20:33:41+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 24.9 | 161 | NA | -9999 | 185 | 399 | 182 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 25.4 | 128 | NA | -9999 | 189 | 449 | 187 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 23.4 | 429 | PASS | -9999 | 260 | 640 | 259 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24.4 | 190 | PASS | -9999 | 104 | 318 | 102 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 33 | NA | -9999 | 144 | 330 | 143 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 21.9 | 421 | PASS | -9999 | 136 | 350 | 133 | PASS | 0 | 40 | QC1 | Probe | 111 | 111 | 111 | PASS | QC2 | Probe | 98 | 97 | 97 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.4 | 60.1 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70.1 | 62.5 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.7 | 175.5 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 18.9 | -9.5 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_09_1_12X | 614386394 | 92303 | Positive Control 1 | 839032 | 210019897 | 6.8 | 2025-10-26T20:34:06+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 25.3 | 132 | NA | -9999 | 202 | 435 | 194 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 26.7 | 111 | NA | -9999 | 174 | 410 | 171 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 24.2 | 395 | PASS | -9999 | 226 | 554 | 223 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 25.1 | 169 | PASS | -9999 | 111 | 336 | 109 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 28 | NA | -9999 | 184 | 432 | 183 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 22.3 | 389 | PASS | -9999 | 141 | 364 | 138 | PASS | 0 | 40 | QC1 | Probe | 117 | 116 | 116 | PASS | QC2 | Probe | 108 | 107 | 107 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.2 | 56.2 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 69.9 | 55.3 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.4 | 169.9 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 19.6 | -11 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_09_1_13X | 614385714 | 92303 | Positive Control 1 | 839032 | 730217 | 6.8 | 2025-10-26T20:34:32+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 24.2 | 144 | NA | -9999 | 233 | 480 | 225 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 26 | 111 | NA | -9999 | 195 | 460 | 193 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 22.9 | 412 | PASS | -9999 | 255 | 618 | 252 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 23.9 | 190 | PASS | -9999 | 92 | 310 | 89 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 23 | NA | -9999 | 147 | 324 | 145 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 21.8 | 397 | PASS | -9999 | 140 | 370 | 138 | PASS | 0 | 40 | QC1 | Probe | 116 | 116 | 115 | PASS | QC2 | Probe | 102 | 101 | 101 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.2 | 57.9 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 69.8 | 54.2 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.4 | 164.6 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 19 | -9.9 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_09_1_14X | 614386482 | 92303 | Positive Control 1 | 839032 | 690802 | 6.8 | 2025-10-26T20:34:58+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 24.9 | 135 | NA | -9999 | 230 | 475 | 226 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 26.5 | 105 | NA | -9999 | 213 | 498 | 211 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 22.9 | 431 | PASS | -9999 | 267 | 648 | 266 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24.3 | 181 | PASS | -9999 | 114 | 343 | 112 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 34 | NA | -9999 | 169 | 380 | 169 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 21.7 | 420 | PASS | -9999 | 143 | 377 | 141 | PASS | 0 | 40 | QC1 | Probe | 116 | 116 | 115 | PASS | QC2 | Probe | 108 | 108 | 108 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.1 | 60.7 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 69.7 | 55.4 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.3 | 177.2 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 17.8 | -9.6 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_09_1_15+ | 614385059 | 92303 | Positive Control 1 | 839032 | 210026044 | 6.8 | 2025-10-26T20:35:22+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 23.3 | 157 | NA | -9999 | 198 | 429 | 189 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 25.3 | 118 | NA | -9999 | 147 | 356 | 144 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 23.1 | 417 | PASS | -9999 | 235 | 587 | 233 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 23.5 | 196 | PASS | -9999 | 79 | 246 | 77 | PASS | 0 | 40 | rpoB3 | Target | 3 | POS | 39.5 | 41 | PASS | -9999 | 179 | 408 | 178 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 20.6 | 486 | PASS | -9999 | 83 | 223 | 81 | PASS | 0 | 40 | QC1 | Probe | 78 | 78 | 77 | PASS | QC2 | Probe | 112 | 112 | 111 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.2 | 57.8 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 69.8 | 44.6 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.4 | 111.4 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 17.5 | -7.9 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_09_1_16+ | 614385485 | 92303 | Positive Control 1 | 839032 | 638845 | 6.8 | 2025-10-26T20:35:47+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 24.6 | 165 | NA | -9999 | 239 | 521 | 231 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 26.9 | 137 | NA | -9999 | 224 | 534 | 221 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 24.7 | 436 | PASS | -9999 | 300 | 751 | 296 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 26 | 178 | PASS | -9999 | 115 | 351 | 114 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 33 | NA | -9999 | 193 | 476 | 192 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 23.7 | 430 | PASS | -9999 | 155 | 408 | 153 | PASS | 0 | 40 | QC1 | Probe | 123 | 123 | 122 | PASS | QC2 | Probe | 113 | 112 | 113 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.4 | 46.7 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70.1 | 38.7 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.6 | 181.5 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 19.5 | -10.7 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_09_1_17+ | 614386401 | 92303 | Positive Control 1 | 839032 | 210035802 | 6.8 | 2025-10-26T20:36:13+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 23 | 170 | NA | -9999 | 231 | 484 | 221 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 25.7 | 124 | NA | -9999 | 191 | 450 | 188 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 23.5 | 440 | PASS | -9999 | 248 | 608 | 245 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24.6 | 190 | PASS | -9999 | 109 | 343 | 107 | PASS | 0 | 40 | rpoB3 | Target | 3 | POS | 35.2 | 48 | PASS | -9999 | 180 | 432 | 179 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 22.4 | 418 | PASS | -9999 | 141 | 367 | 139 | PASS | 0 | 40 | QC1 | Probe | 116 | 116 | 115 | PASS | QC2 | Probe | 118 | 118 | 118 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.3 | 47.7 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70 | 47.3 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.6 | 162.3 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 19.9 | -8.9 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_09_1_18+ | 614385629 | 92303 | Positive Control 1 | 839032 | 880236 | 6.8 | 2025-10-26T20:36:38+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 24.1 | 169 | NA | -9999 | 189 | 409 | 178 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 26 | 135 | NA | -9999 | 188 | 446 | 184 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 23.3 | 457 | PASS | -9999 | 231 | 572 | 226 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24.5 | 188 | PASS | -9999 | 109 | 328 | 106 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 34 | NA | -9999 | 156 | 371 | 154 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 22.2 | 406 | PASS | -9999 | 136 | 351 | 132 | PASS | 0 | 40 | QC1 | Probe | 118 | 117 | 117 | PASS | QC2 | Probe | 100 | 100 | 100 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.3 | 43.4 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70 | 47.4 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.5 | 147.8 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 20.2 | -8.2 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_10_1_11X | 614388914 | 92303 | Positive Control 1 | 839032 | 210016126 | 6.8 | 2025-10-26T22:34:44+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 24.5 | 157 | NA | -9999 | 237 | 504 | 235 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 25.6 | 130 | NA | -9999 | 202 | 482 | 201 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 23.1 | 456 | PASS | -9999 | 288 | 717 | 287 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24.1 | 194 | PASS | -9999 | 112 | 339 | 111 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 38 | NA | -9999 | 174 | 419 | 174 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 21.7 | 442 | PASS | -9999 | 143 | 372 | 141 | PASS | 0 | 40 | QC1 | Probe | 116 | 117 | 116 | PASS | QC2 | Probe | 99 | 99 | 99 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.3 | 58.8 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70 | 54.1 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.5 | 173.7 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 15.3 | -8 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_10_1_12X | 614389210 | 92303 | Positive Control 1 | 839032 | 736872 | 6.8 | 2025-10-26T22:35:10+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 24.7 | 160 | NA | -9999 | 181 | 409 | 175 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 26.8 | 121 | NA | -9999 | 210 | 495 | 207 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 24.6 | 413 | PASS | -9999 | 264 | 658 | 262 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 25.4 | 184 | PASS | -9999 | 105 | 331 | 103 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 29 | NA | -9999 | 180 | 431 | 180 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 23.1 | 408 | PASS | -9999 | 139 | 372 | 136 | PASS | 0 | 40 | QC1 | Probe | 106 | 106 | 105 | PASS | QC2 | Probe | 110 | 110 | 110 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.4 | 53.3 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70 | 49.1 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.6 | 174.8 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 20.9 | -10.6 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_10_1_13X | 614389594 | 92303 | Positive Control 1 | 839032 | 739871 | 6.8 | 2025-10-26T22:35:33+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 24.7 | 159 | NA | -9999 | 203 | 450 | 197 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 25.4 | 129 | NA | -9999 | 198 | 471 | 195 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 23.2 | 442 | PASS | -9999 | 273 | 675 | 271 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24.4 | 186 | PASS | -9999 | 112 | 342 | 110 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 34 | NA | -9999 | 167 | 406 | 166 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 22 | 426 | PASS | -9999 | 143 | 387 | 140 | PASS | 0 | 40 | QC1 | Probe | 115 | 115 | 114 | PASS | QC2 | Probe | 107 | 107 | 107 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.4 | 53.6 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70.1 | 48.8 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.6 | 174.6 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 18.4 | -9.2 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_10_1_14X | 614388652 | 92303 | Positive Control 1 | 839032 | 656159 | 6.8 | 2025-10-26T22:35:57+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 23.8 | 165 | NA | -9999 | 189 | 409 | 184 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 26 | 126 | NA | -9999 | 218 | 519 | 213 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 23.2 | 455 | PASS | -9999 | 275 | 689 | 272 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24.4 | 190 | PASS | -9999 | 117 | 353 | 114 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 31 | NA | -9999 | 173 | 413 | 172 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 22.2 | 432 | PASS | -9999 | 137 | 359 | 135 | PASS | 0 | 40 | QC1 | Probe | 111 | 110 | 110 | PASS | QC2 | Probe | 111 | 111 | 110 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.4 | 50.9 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70.1 | 47.7 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.6 | 158 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 17.7 | -11.9 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_10_1_15+ | 614389737 | 92303 | Positive Control 1 | 839032 | 210059038 | 6.8 | 2025-10-26T22:36:21+0100 | Done | MTB DETECTED LOW;RIF Resistance INDETERMINATE | SPC | SPC | 1,2,3,4 | NA | 24.2 | 160 | NA | -9999 | 145 | 331 | 141 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 26 | 118 | NA | -9999 | 196 | 464 | 193 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 23.5 | 423 | PASS | -9999 | 239 | 596 | 235 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24.2 | 192 | PASS | -9999 | 97 | 300 | 95 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 35 | NA | -9999 | 175 | 415 | 174 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 22.1 | 416 | PASS | -9999 | 140 | 363 | 136 | PASS | 0 | 40 | QC1 | Probe | 114 | 113 | 113 | PASS | QC2 | Probe | 109 | 109 | 109 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.4 | 46 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70 | 42.2 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.5 | 159.9 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 16.4 | -7.6 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_10_1_16+ | 614389057 | 92303 | Positive Control 1 | 839032 | 210012379 | 6.8 | 2025-10-26T22:36:46+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 24.4 | 143 | NA | -9999 | 191 | 418 | 186 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 25.9 | 123 | NA | -9999 | 200 | 477 | 197 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 23.6 | 409 | PASS | -9999 | 235 | 578 | 232 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24.4 | 182 | PASS | -9999 | 108 | 326 | 106 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 32 | NA | -9999 | 182 | 430 | 181 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 22.2 | 400 | PASS | -9999 | 131 | 341 | 129 | PASS | 0 | 40 | QC1 | Probe | 105 | 104 | 104 | PASS | QC2 | Probe | 120 | 119 | 120 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.3 | 51.5 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70 | 49.9 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.5 | 155.2 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 18.3 | -7.4 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_10_1_17+ | 614389329 | 92303 | Positive Control 1 | 839032 | 210003146 | 6.8 | 2025-10-26T22:37:10+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 24.2 | 151 | NA | -9999 | 164 | 352 | 159 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 25.4 | 120 | NA | -9999 | 163 | 391 | 161 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 23.4 | 410 | PASS | -9999 | 238 | 586 | 238 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24 | 188 | PASS | -9999 | 93 | 294 | 91 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 27 | NA | -9999 | 150 | 357 | 150 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 21.6 | 417 | PASS | -9999 | 125 | 328 | 123 | PASS | 0 | 40 | QC1 | Probe | 95 | 95 | 95 | PASS | QC2 | Probe | 89 | 89 | 89 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.2 | 59.9 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 69.8 | 52.9 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.3 | 149 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 20.5 | -8.4 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_10_1_18+ | 614388817 | 92303 | Positive Control 1 | 839032 | 694834 | 6.8 | 2025-10-26T22:37:37+0100 | Done | MTB DETECTED LOW;RIF Resistance DETECTED | SPC | SPC | 1,2,3,4 | NA | 24.5 | 142 | NA | -9999 | 206 | 439 | 199 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 25.5 | 111 | NA | -9999 | 197 | 468 | 193 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 23.2 | 396 | PASS | -9999 | 262 | 641 | 259 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 23.9 | 184 | PASS | -9999 | 118 | 366 | 116 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 30 | NA | -9999 | 169 | 399 | 168 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 21.5 | 414 | PASS | -9999 | 160 | 424 | 157 | PASS | 0 | 40 | QC1 | Probe | 123 | 123 | 122 | PASS | QC2 | Probe | 110 | 110 | 110 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.3 | 53.3 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 69.9 | 51.5 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.5 | 166.4 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 17.8 | -9.4 | [5 empty] | <None>
Enter Resample Data Below: |  | ¨N/A
Assay | Assay Version | Sample ID | Cartridge S/N | Reagent Lot ID | Test Type | Instrument S/N | Module S/N | S/W Version | Start Time | Status | Test Result | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Max Pressure (PSI) | Min Pressure (PSI) | Error | Patient ID | Quantitative Value | Quantitative Unit | Interpretation Flag | User

## Data - Sample Type 2
Assay | Assay Version | Sample ID | Cartridge S/N | Reagent Lot ID | Test Type | Instrument S/N | Module S/N | S/W Version | Start Time | Status | Test Result | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Max Pressure (PSI) | Min Pressure (PSI) | Error | Patient ID | Quantitative Value | Quantitative Unit | Interpretation Flag | User
Xpert MTB-RIF Ultra | 4 | B01NOV24B_01_2_19X | 614332206 | 92303 | Positive Control 2 | 802069 | 656295 | 7.1.6 | 2025-10-26T20:32:59+0100 | Done | MTB DETECTED LOW;RIF Resistance NOT DETECTED | SPC | SPC | 1,2,3,4 | NA | 24.1 | 160 | NA | -9999 | 205 | 432 | 199 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 22.4 | 453 | NA | -9999 | 169 | 402 | 167 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 24.7 | 407 | PASS | -9999 | 229 | 562 | 225 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 25 | 253 | PASS | -9999 | 99 | 296 | 97 | PASS | 0 | 40 | rpoB3 | Target | 3 | POS | 25.8 | 188 | PASS | -9999 | 153 | 361 | 152 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 27.2 | 138 | PASS | -9999 | 128 | 329 | 126 | PASS | 0 | 40 | QC1 | Probe | 105 | 105 | 105 | PASS | QC2 | Probe | 97 | 96 | 96 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69 | 68.6 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA | 72.8 | 98 | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA | 75.5 | 88.3 | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA | 66.8 | 93.1 | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 15.6 | -8.1 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | B01NOV24B_01_2_20+ | 614332203 | 92303 | Positive Control 2 | 802069 | 624207 | 7.1.6 | 2025-10-26T20:36:56+0100 | Done | MTB DETECTED LOW;RIF Resistance NOT DETECTED | SPC | SPC | 1,2,3,4 | NA | 24.4 | 161 | NA | -9999 | 189 | 406 | 181 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 21.1 | 472 | NA | -9999 | 174 | 412 | 171 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 23.6 | 442 | PASS | -9999 | 233 | 578 | 231 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 23.6 | 285 | PASS | -9999 | 95 | 288 | 92 | PASS | 0 | 40 | rpoB3 | Target | 3 | POS | 24.7 | 205 | PASS | -9999 | 170 | 404 | 169 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 25.8 | 143 | PASS | -9999 | 129 | 342 | 126 | PASS | 0 | 40 | QC1 | Probe | 98 | 98 | 97 | PASS | QC2 | Probe | 95 | 95 | 95 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69 | 78.8 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA | 72.8 | 98.2 | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA | 75.5 | 92.9 | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA | 66.9 | 93.8 | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 20.8 | -8.4 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | B01NOV24B_02_2_19X | 614334366 | 92303 | Positive Control 2 | 802069 | 642168 | 7.1.6 | 2025-10-26T20:33:23+0100 | Done | MTB DETECTED LOW;RIF Resistance NOT DETECTED | SPC | SPC | 1,2,3,4 | NA | 24.9 | 160 | NA | -9999 | 214 | 460 | 206 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 23.1 | 483 | NA | -9999 | 154 | 362 | 152 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 25 | 441 | PASS | -9999 | 256 | 633 | 250 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 25.1 | 285 | PASS | -9999 | 90 | 284 | 89 | PASS | 0 | 40 | rpoB3 | Target | 3 | POS | 26.4 | 191 | PASS | -9999 | 162 | 397 | 162 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 27.4 | 148 | PASS | -9999 | 123 | 324 | 119 | PASS | 0 | 40 | QC1 | Probe | 96 | 95 | 95 | PASS | QC2 | Probe | 103 | 103 | 103 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.3 | 68.3 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA | 73 | 92.3 | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA | 75.8 | 89.1 | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA | 67 | 98.9 | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 18.7 | -8.2 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | B01NOV24B_02_2_20+ | 614334511 | 92303 | Positive Control 2 | 802069 | 775844 | 7.1.6 | 2025-10-26T20:37:22+0100 | Done | MTB DETECTED LOW;RIF Resistance NOT DETECTED | SPC | SPC | 1,2,3,4 | NA | 25.1 | 137 | NA | -9999 | 191 | 411 | 187 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 23.8 | 484 | NA | -9999 | 163 | 389 | 161 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 26.1 | 379 | PASS | -9999 | 232 | 563 | 230 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 26.3 | 242 | PASS | -9999 | 91 | 301 | 89 | PASS | 0 | 40 | rpoB3 | Target | 3 | POS | 27.8 | 154 | PASS | -9999 | 175 | 414 | 175 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 28.7 | 119 | PASS | -9999 | 152 | 397 | 150 | PASS | 0 | 40 | QC1 | Probe | 108 | 108 | 108 | PASS | QC2 | Probe | 99 | 99 | 99 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69 | 71.8 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA | 72.7 | 98.3 | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA | 75.4 | 86.9 | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA | 66.6 | 103.6 | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 21.9 | -12.3 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | B01NOV24B_03_2_19X | 614337470 | 92303 | Positive Control 2 | 802069 | 635193 | 7.1.6 | 2025-10-26T20:33:47+0100 | Aborted | NO RESULT | SPC | SPC | 1,2,3,4 | NA | 24.6 | 154 | NA | -9999 | 208 | 438 | 203 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 23.6 | 481 | NA | -9999 | 209 | 499 | 207 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 25.2 | 427 | PASS | -9999 | 306 | 761 | 303 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 25.2 | 273 | PASS | -9999 | 116 | 353 | 115 | PASS | 0 | 40 | rpoB3 | Target | 3 | POS | 26.3 | 190 | PASS | -9999 | 154 | 374 | 155 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 27.9 | 132 | PASS | -9999 | 142 | 372 | 140 | PASS | 0 | 40 | QC1 | Probe | 113 | 113 | 113 | PASS | QC2 | Probe | 93 | 92 | 92 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.2 | 81.3 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA | 72.9 | 106.2 | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA | 75.7 | 93.6 | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA | 66.9 | 107.8 | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 20 | -8.9 | Error 2037: The cartridge integrity test failed at valve position 1390. The pressure change of -0.6 PSI did not exceed the requirement of 4.0 PSI. The pressure increased from 3.5 PSI to 3.0 PSI during the test |  |  |  |  | <None>
Xpert MTB-RIF Ultra | 4 | B01NOV24B_03_2_20+ | 614337317 | 92303 | Positive Control 2 | 802069 | 786669 | 7.1.6 | 2025-10-26T20:37:47+0100 | Done | MTB DETECTED LOW;RIF Resistance NOT DETECTED | SPC | SPC | 1,2,3,4 | NA | 24.4 | 167 | NA | -9999 | 185 | 405 | 177 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 24.4 | 456 | NA | -9999 | 183 | 434 | 180 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 25.2 | 423 | PASS | -9999 | 245 | 611 | 242 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 25.2 | 271 | PASS | -9999 | 109 | 328 | 106 | PASS | 0 | 40 | rpoB3 | Target | 3 | POS | 26.1 | 192 | PASS | -9999 | 158 | 381 | 157 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 27.4 | 134 | PASS | -9999 | 127 | 328 | 124 | PASS | 0 | 40 | QC1 | Probe | 97 | 97 | 96 | PASS | QC2 | Probe | 95 | 94 | 94 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.2 | 67.3 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA | 73 | 92.9 | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA | 75.7 | 85.2 | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA | 67 | 77.7 | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 20.8 | -15 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | B01NOV24B_04_2_19X | 614341778 | 92303 | Positive Control 2 | 802069 | 659869 | 7.1.6 | 2025-10-26T20:34:11+0100 | Done | MTB DETECTED LOW;RIF Resistance NOT DETECTED | SPC | SPC | 1,2,3,4 | NA | 0 | 144 | NA | -9999 | 235 | 499 | 227 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 23.2 | 467 | NA | -9999 | 171 | 406 | 170 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 24.5 | 420 | PASS | -9999 | 245 | 603 | 243 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24.3 | 277 | PASS | -9999 | 107 | 322 | 105 | PASS | 0 | 40 | rpoB3 | Target | 3 | POS | 25.4 | 191 | PASS | -9999 | 159 | 370 | 159 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 26.8 | 130 | PASS | -9999 | 139 | 361 | 137 | PASS | 0 | 40 | QC1 | Probe | 111 | 110 | 110 | PASS | QC2 | Probe | 94 | 94 | 94 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69 | 75.2 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA | 72.7 | 98.9 | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA | 75.4 | 83.3 | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA | 66.8 | 85.4 | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 19.7 | -6.9 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | B01NOV24B_04_2_20+ | 614341609 | 92303 | Positive Control 2 | 802069 | 210052215 | 7.1.6 | 2025-10-26T20:38:11+0100 | Done | MTB DETECTED LOW;RIF Resistance NOT DETECTED | SPC | SPC | 1,2,3,4 | NA | 26 | 154 | NA | -9999 | 185 | 397 | 178 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 23 | 461 | NA | -9999 | 162 | 387 | 161 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 25.2 | 403 | PASS | -9999 | 225 | 549 | 223 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 25.1 | 265 | PASS | -9999 | 104 | 317 | 102 | PASS | 0 | 40 | rpoB3 | Target | 3 | POS | 26.1 | 183 | PASS | -9999 | 146 | 343 | 146 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 26.6 | 146 | PASS | -9999 | 124 | 321 | 123 | PASS | 0 | 40 | QC1 | Probe | 107 | 106 | 106 | PASS | QC2 | Probe | 82 | 82 | 82 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.1 | 66.2 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA | 72.8 | 94.8 | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA | 75.6 | 76.1 | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA | 66.8 | 79 | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 20.9 | -9 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | B01NOV24B_05_2_19X | 614344026 | 92303 | Positive Control 2 | 802069 | 903819 | 7.1.6 | 2025-10-26T20:34:37+0100 | Done | MTB DETECTED LOW;RIF Resistance NOT DETECTED | SPC | SPC | 1,2,3,4 | NA | 24 | 138 | NA | -9999 | 228 | 481 | 220 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 22.7 | 456 | NA | -9999 | 191 | 449 | 186 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 24.3 | 421 | PASS | -9999 | 257 | 629 | 251 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24.5 | 267 | PASS | -9999 | 114 | 348 | 110 | PASS | 0 | 40 | rpoB3 | Target | 3 | POS | 25.8 | 179 | PASS | -9999 | 165 | 392 | 163 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 26.7 | 137 | PASS | -9999 | 124 | 321 | 121 | PASS | 0 | 40 | QC1 | Probe | 108 | 107 | 107 | PASS | QC2 | Probe | 105 | 105 | 105 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.2 | 73.1 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA | 73 | 102.9 | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA | 75.7 | 84.4 | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA | 67 | 85.6 | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 18.3 | -7.3 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | B01NOV24B_05_2_20+ | 614345151 | 92303 | Positive Control 2 | 802069 | 805086 | 7.1.6 | 2025-10-26T20:38:38+0100 | Done | MTB DETECTED LOW;RIF Resistance NOT DETECTED | SPC | SPC | 1,2,3,4 | NA | 24.7 | 164 | NA | -9999 | 173 | 379 | 166 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 23 | 466 | NA | -9999 | 166 | 395 | 164 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 25.1 | 410 | PASS | -9999 | 234 | 580 | 232 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24.8 | 269 | PASS | -9999 | 108 | 333 | 107 | PASS | 0 | 40 | rpoB3 | Target | 3 | POS | 25.7 | 191 | PASS | -9999 | 159 | 377 | 158 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 27.2 | 135 | PASS | -9999 | 132 | 347 | 130 | PASS | 0 | 40 | QC1 | Probe | 110 | 110 | 109 | PASS | QC2 | Probe | 100 | 100 | 99 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.2 | 66.8 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA | 73 | 98.4 | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA | 75.7 | 84.2 | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA | 66.9 | 88.4 | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 21.3 | -8.7 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | B01NOV24B_06_2_19X | 614348164 | 92303 | Positive Control 2 | 802069 | 621821 | 7.1.6 | 2025-10-26T20:35:02+0100 | Done | MTB DETECTED LOW;RIF Resistance NOT DETECTED | SPC | SPC | 1,2,3,4 | NA | 23.6 | 155 | NA | -9999 | 173 | 379 | 167 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 24 | 455 | NA | -9999 | 165 | 392 | 162 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 25.4 | 396 | PASS | -9999 | 251 | 618 | 249 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 25.3 | 264 | PASS | -9999 | 102 | 320 | 99 | PASS | 0 | 40 | rpoB3 | Target | 3 | POS | 26.3 | 184 | PASS | -9999 | 169 | 404 | 168 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 26.8 | 162 | PASS | -9999 | 126 | 332 | 123 | PASS | 0 | 40 | QC1 | Probe | 98 | 97 | 96 | PASS | QC2 | Probe | 110 | 110 | 110 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.3 | 55.1 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA | 72.9 | 39.9 | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA | 75.7 | 52.7 | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA | 66.8 | 50.5 | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 26.5 | -13.5 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | 06_2_20+ | 0 | 92303 | [6 empty] | V
Xpert MTB-RIF Ultra | 4 | B01NOV24B_07_2_19X | 614348958 | 92303 | Positive Control 2 | 802069 | 884165 | 7.1.6 | 2025-10-26T20:35:27+0100 | Done | MTB DETECTED LOW;RIF Resistance NOT DETECTED | SPC | SPC | 1,2,3,4 | NA | 24.6 | 162 | NA | -9999 | 166 | 362 | 159 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 23 | 477 | NA | -9999 | 165 | 393 | 162 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 24.6 | 422 | PASS | -9999 | 240 | 596 | 237 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24.7 | 262 | PASS | -9999 | 96 | 293 | 94 | PASS | 0 | 40 | rpoB3 | Target | 3 | POS | 25.8 | 183 | PASS | -9999 | 153 | 345 | 152 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 27.4 | 129 | PASS | -9999 | 122 | 313 | 119 | PASS | 0 | 40 | QC1 | Probe | 98 | 98 | 98 | PASS | QC2 | Probe | 96 | 95 | 95 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.2 | 72.9 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA | 73 | 86.9 | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA | 75.7 | 82.9 | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA | 67 | 81.4 | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 25.6 | -12.3 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | B01NOV24B_07_2_20+ | 614349611 | 92303 | Positive Control 2 | 802069 | 869520 | 7.1.6 | 2025-10-26T20:39:04+0100 | Done | MTB DETECTED LOW;RIF Resistance NOT DETECTED | SPC | SPC | 1,2,3,4 | NA | 23.4 | 168 | NA | -9999 | 138 | 304 | 136 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 23.1 | 484 | NA | -9999 | 181 | 432 | 181 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 25.1 | 432 | PASS | -9999 | 211 | 531 | 211 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 25 | 284 | PASS | -9999 | 99 | 315 | 97 | PASS | 0 | 40 | rpoB3 | Target | 3 | POS | 26.2 | 191 | PASS | -9999 | 159 | 376 | 159 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 27.6 | 135 | PASS | -9999 | 130 | 344 | 129 | PASS | 0 | 40 | QC1 | Probe | 105 | 104 | 104 | PASS | QC2 | Probe | 93 | 93 | 92 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.1 | 63.6 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA | 72.9 | 97.8 | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA | 75.6 | 85.6 | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA | 66.9 | 88.9 | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 30 | -16.5 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | B01NOV24B_08_2_19X | 614383062 | 92303 | Positive Control 2 | 802069 | 656453 | 7.1.6 | 2025-10-26T20:35:53+0100 | Done | MTB DETECTED LOW;RIF Resistance NOT DETECTED | SPC | SPC | 1,2,3,4 | NA | 25.3 | 133 | NA | -9999 | 177 | 378 | 173 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 23.7 | 479 | NA | -9999 | 166 | 393 | 164 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 26.8 | 364 | PASS | -9999 | 222 | 547 | 220 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 26.9 | 241 | PASS | -9999 | 89 | 284 | 87 | PASS | 0 | 40 | rpoB3 | Target | 3 | POS | 28.1 | 164 | PASS | -9999 | 145 | 339 | 144 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 29.3 | 131 | PASS | -9999 | 111 | 293 | 108 | PASS | 0 | 40 | QC1 | Probe | 85 | 85 | 84 | PASS | QC2 | Probe | 97 | 96 | 96 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.1 | 59.7 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA | 72.9 | 83 | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA | 75.6 | 76.3 | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA | 66.9 | 81.4 | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 20 | -10.5 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | B01NOV24B_08_2_20+ | 614382995 | 92303 | Positive Control 2 | 802069 | 719652 | 7.1.6 | 2025-10-26T20:39:28+0100 | Done | MTB DETECTED LOW;RIF Resistance NOT DETECTED | SPC | SPC | 1,2,3,4 | NA | 23.4 | 145 | NA | -9999 | 213 | 444 | 206 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 22.6 | 465 | NA | -9999 | 184 | 442 | 182 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 24.6 | 380 | PASS | -9999 | 251 | 619 | 250 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 23.9 | 269 | PASS | -9999 | 98 | 303 | 96 | PASS | 0 | 40 | rpoB3 | Target | 3 | POS | 25.1 | 186 | PASS | -9999 | 161 | 367 | 160 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 26.7 | 125 | PASS | -9999 | 142 | 373 | 140 | PASS | 0 | 40 | QC1 | Probe | 119 | 119 | 118 | PASS | QC2 | Probe | 109 | 109 | 108 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.1 | 72.4 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA | 72.7 | 80.9 | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA | 75.4 | 74.1 | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA | 66.7 | 75.5 | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 23.7 | -13.4 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | B01NOV24B_09_2_19X | 614385874 | 92303 | Positive Control 2 | 802069 | 679144 | 7.1.6 | 2025-10-26T20:36:29+0100 | Done | MTB DETECTED LOW;RIF Resistance NOT DETECTED | SPC | SPC | 1,2,3,4 | NA | 23.5 | 163 | NA | -9999 | 160 | 340 | 153 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 21.9 | 450 | NA | -9999 | 172 | 406 | 167 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 24.1 | 414 | PASS | -9999 | 219 | 534 | 215 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 23.9 | 275 | PASS | -9999 | 100 | 300 | 96 | PASS | 0 | 40 | rpoB3 | Target | 3 | POS | 25.1 | 190 | PASS | -9999 | 157 | 363 | 155 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 25.8 | 141 | PASS | -9999 | 126 | 325 | 122 | PASS | 0 | 40 | QC1 | Probe | 108 | 107 | 107 | PASS | QC2 | Probe | 106 | 106 | 105 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.1 | 65.9 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA | 72.8 | 92.9 | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA | 75.5 | 81.5 | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA | 66.9 | 84.6 | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 20.2 | -6.8 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | B01NOV24B_09_2_20+ | 614386021 | 92303 | Positive Control 2 | 802069 | 686014 | 7.1.6 | 2025-10-26T20:39:52+0100 | Done | MTB DETECTED LOW;RIF Resistance NOT DETECTED | SPC | SPC | 1,2,3,4 | NA | 23.9 | 152 | NA | -9999 | 193 | 427 | 189 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 21.5 | 454 | NA | -9999 | 215 | 511 | 214 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 24.2 | 410 | PASS | -9999 | 287 | 718 | 286 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 23.7 | 289 | PASS | -9999 | 94 | 296 | 92 | PASS | 0 | 40 | rpoB3 | Target | 3 | POS | 25.1 | 191 | PASS | -9999 | 163 | 398 | 163 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 26.5 | 136 | PASS | -9999 | 135 | 354 | 133 | PASS | 0 | 40 | QC1 | Probe | 110 | 110 | 110 | PASS | QC2 | Probe | 115 | 115 | 115 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.2 | 80.3 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA | 72.9 | 84.1 | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA | 75.7 | 84.1 | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA | 67 | 81 | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 20.5 | -11.1 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | B01NOV24B_10_2_19X | 614388340 | 92303 | Positive Control 2 | 839032 | 681294 | 6.8 | 2025-10-26T22:38:02+0100 | Done | MTB DETECTED LOW;RIF Resistance NOT DETECTED | SPC | SPC | 1,2,3,4 | NA | 25.4 | 143 | NA | -9999 | 192 | 418 | 186 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 21.8 | 432 | NA | -9999 | 196 | 465 | 194 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 24.8 | 360 | PASS | -9999 | 281 | 691 | 278 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 23.2 | 300 | PASS | -9999 | 129 | 395 | 127 | PASS | 0 | 40 | rpoB3 | Target | 3 | POS | 25.3 | 183 | PASS | -9999 | 163 | 395 | 162 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 26.5 | 127 | PASS | -9999 | 130 | 341 | 128 | PASS | 0 | 40 | QC1 | Probe | 107 | 107 | 107 | PASS | QC2 | Probe | 99 | 98 | 98 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.1 | 91.9 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA | 72.9 | 143.9 | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA | 75.6 | 95.2 | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA | 67 | 86.4 | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 15.4 | -13.2 | [5 empty] | <None>
Xpert MTB-RIF Ultra | 4 | B01NOV24B_10_2_20+ | 614388439 | 92303 | Positive Control 2 | 839032 | 210058622 | 6.8 | 2025-10-26T22:38:28+0100 | Done | MTB DETECTED LOW;RIF Resistance NOT DETECTED | SPC | SPC | 1,2,3,4 | NA | 24 | 185 | NA | -9999 | 152 | 335 | 146 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 21.9 | 480 | NA | -9999 | 195 | 464 | 193 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 23.1 | 494 | PASS | -9999 | 250 | 623 | 249 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 23.3 | 305 | PASS | -9999 | 116 | 357 | 114 | PASS | 0 | 40 | rpoB3 | Target | 3 | POS | 24.2 | 224 | PASS | -9999 | 164 | 395 | 164 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 25.4 | 159 | PASS | -9999 | 158 | 413 | 155 | PASS | 0 | 40 | QC1 | Probe | 126 | 126 | 126 | PASS | QC2 | Probe | 109 | 109 | 109 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.2 | 70.8 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA | 73.1 | 110.1 | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA | 75.8 | 95.4 | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA | 67.1 | 114.4 | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 17.5 | -7.1 | [5 empty] | <None>
Enter Resample Data Below: |  | ¨N/A
Assay | Assay Version | Sample ID | Cartridge S/N | Reagent Lot ID | Test Type | Instrument S/N | Module S/N | S/W Version | Start Time | Status | Test Result | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Max Pressure (PSI) | Min Pressure (PSI) | Error | Patient ID | Quantitative Value | Quantitative Unit | Interpretation Flag | User

## Data - Replaced
Assay | Assay Version | Sample ID | Cartridge S/N | Reagent Lot ID | Test Type | Instrument S/N | Module S/N | S/W Version | Start Time | Status | Test Result | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Max Pressure (PSI) | Min Pressure (PSI) | Error | Patient ID | Quantitative Value | Quantitative Unit | Interpretation Flag | User
Xpert MTB-RIF Ultra | 4 | 01_0_08+ | 0 | 92303 | [6 empty] | S | [170 empty] | Seal Test Failure
Xpert MTB-RIF Ultra | 4 | 04_1_14X | 0 | 92303 | [6 empty] | S
Xpert MTB-RIF Ultra | 4 | C15NOV23D_03_1_14X | 614338042 | 92303 | Positive Control 1 | 839032 | 656159 | 6.8 | 2025-10-26T04:57:00+0100 | Aborted | ERROR | SPC | SPC | 1,2,3,4 | NA | 24.4 | 154 | NA | -9999 | 181 | 393 | 177 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 25.9 | 121 | NA | -9999 | 207 | 494 | 204 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 23.7 | 410 | PASS | -9999 | 262 | 648 | 258 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24.8 | 177 | PASS | -9999 | 114 | 344 | 111 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 31 | NA | -9999 | 163 | 391 | 163 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 22.4 | 411 | PASS | -9999 | 131 | 341 | 128 | PASS | 0 | 40 | QC1 | Probe | 106 | 105 | 105 | PASS | QC2 | Probe | 108 | 108 | 108 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.4 | 55.4 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70 | 48.9 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.5 | 152 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 19.2 | -10 | Error 2005: Motion of the syringe drive was not detected. Detected motion started at position 250 uL and transferred 211 uL at valve position 245 with pressure 2.7 PSI |  |  |  |  | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_03_1_15+ | 614338040 | 92303 | Positive Control 1 | 839032 | 210059038 | 6.8 | 2025-10-26T04:57:26+0100 | Done | ERROR | SPC | SPC | 1,2,3,4 | NA | 24 | 155 | NA | -9999 | 146 | 344 | 142 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 27.5 | 125 | NA | -9999 | 202 | 478 | 199 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 25.4 | 393 | PASS | -9999 | 253 | 634 | 250 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 26 | 182 | PASS | -9999 | 101 | 309 | 99 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 32 | NA | -9999 | 182 | 437 | 181 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 24.1 | 392 | PASS | -9999 | 131 | 347 | 129 | PASS | 0 | 40 | QC1 | Probe | 115 | 114 | 114 | PASS | QC2 | Probe | 122 | 122 | 122 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.4 | 41.2 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70.1 | 36.4 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.6 | 139.7 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 24.6 | -10 | Error 1002: The temperature difference of 5.3 ºC exceeds the limit of 5.0 ºC. The temperatures for heaters A and B are 97.3 ºC and 92.0 ºC |  |  |  |  | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_03_1_16+ | 614338465 | 92303 | Positive Control 1 | 839032 | 210012379 | 6.8 | 2025-10-26T04:57:53+0100 | Done | ERROR | SPC | SPC | 1,2,3,4 | NA | 23.4 | 160 | NA | -9999 | 192 | 422 | 186 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 26 | 130 | NA | -9999 | 207 | 491 | 203 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 23.4 | 435 | PASS | -9999 | 241 | 597 | 237 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 24.4 | 188 | PASS | -9999 | 111 | 332 | 109 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 37 | NA | -9999 | 186 | 447 | 184 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 22.2 | 430 | PASS | -9999 | 133 | 343 | 129 | PASS | 0 | 40 | QC1 | Probe | 114 | 113 | 113 | PASS | QC2 | Probe | 129 | 129 | 129 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.4 | 44.7 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 70.1 | 39.1 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.6 | 138.3 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 17.3 | -6.9 | Error 1001: The actual temperature 61.6 ºC has drifted too far away from the setpoint of 61.0 ºC¦Error 1001: The actual temperature 60.5 ºC has drifted too far away from the setpoint of 61.0 ºC¦Error 1001: The actual temperature 60.4 ºC has drifted too far away from the setpoint of 61.0 ºC¦Error 1001: The actual temperature 61.7 ºC has drifted too far away from the setpoint of 61.0 ºC |  |  |  |  | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_06_1_13AX | 614345586 | 92303 | Positive Control 1 | 839032 | 210018949 | 6.8 | 2025-10-26T13:14:16+0100 | Aborted | ERROR | SPC | SPC | 1,2,3,4 | NA | 23.5 | 161 | NA | -9999 | 240 | 499 | 234 | PASS | 0 | 40 | IS1081-IS6110 | IC | 1,2,3,4 | NA | 25.7 | 117 | NA | -9999 | 232 | 540 | 229 | PASS | 0 | 40 | rpoB1 | Target | 1 | POS | 22.3 | 459 | PASS | -9999 | 299 | 736 | 297 | PASS | 0 | 40 | rpoB2 | Target | 2 | POS | 23.5 | 194 | PASS | -9999 | 123 | 369 | 121 | PASS | 0 | 40 | rpoB3 | Target | 3 | NEG | 0 | 37 | NA | -9999 | 189 | 428 | 188 | PASS | 0 | 40 | rpoB4 | Target | 4 | POS | 21.2 | 431 | PASS | -9999 | 160 | 419 | 157 | PASS | 0 | 40 | QC1 | Probe | 136 | 135 | 135 | PASS | QC2 | Probe | 129 | 128 | 128 | PASS | rpoB1 melt | Melt Target | 0 | 0 | 0 | NA | 69.2 | 61.6 | rpoB2 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB3 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB1 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB2 Mut melt | Melt Target | 0 | 0 | 0 | NA | 69.9 | 57.2 | rpoB3 Mut melt | Melt Target | 0 | 0 | 0 | NA |  |  | rpoB4 Mut melt A | Melt Target | 0 | 0 | 0 | NA | 73.5 | 190 | rpoB4 Mut melt B | Melt Target | 0 | 0 | 0 | NA |  |  | 18 | -7.3 | Error 2005: Motion of the syringe drive was not detected. Detected motion started at position 250 uL and transferred 211 uL at valve position 245 with pressure 2.7 PSI |  |  |  |  | <None>
Xpert MTB-RIF Ultra | 4 | C15NOV23D_06_1_13X | 0 | 92303 | [6 empty] | S

## Statistical Process Control
| Historical Defect Rate: | 2% |  | (1) Input "Sample Size of Subgroup" and "Defects in Subgroup".
| Default Sample Size: | 25 |  | (2) Contact the Supervisor immediately if:
| Confidence: | 0.9 |  |  | (a) the Blue Data Line is above the Red Control Line
and/or
(b) the "Defects in Subgroup" value is displayed in Red
Sampling | Sample Size | Defects |  |  |  |  | Sample Size | Defect
Point | of Subgroup | in Subgroup |  |  |  |  | Per Sub Group | Per Sub Group
1 | 20 | 0 | 4 |  |  | 1 | 20 | 0
2 | 20 | 0 | 4 |  |  | 2 | 20 | 0
3 | 20 | 1 | 4 |  |  | 3 | 20 | 1
4 | 20 | 0 | 4 |  |  | 4 | 20 | 0
5 | 20 | 0 | 4 |  |  | 5 | 20 | 0
6 | 20 | 1 | 4 |  |  | 6 | 20 | 1
7 | 20 | 0 | 4 |  |  | 7 | 20 | 0
8 | 20 | 0 | 4 |  |  | 8 | 20 | 0
9 | 20 | 0 | 4 |  |  | 9 | 20 | 0
10 | 20 | 0 | 4 |  |  | 10 | 20 | 0
11 | N/A | N/A | #VALUE! |  |  | 11 | N/A | N/A
12 | N/A | N/A | #VALUE! |  |  | 12 | N/A | N/A
13 | N/A | N/A | #VALUE! |  |  | 13 | N/A | N/A
14 | N/A | N/A | #VALUE! |  |  | 14 | N/A | N/A
15 | N/A | N/A | #VALUE! |  |  | 15 | N/A | N/A
16 | N/A | N/A | #VALUE! |  |  | 16 | N/A | N/A
17 | N/A | N/A | #VALUE! |  |  | 17 | N/A | N/A
18 | N/A | N/A | #VALUE! |  |  | 18 | N/A | N/A
19 | N/A | N/A | #VALUE! |  |  | 19 | N/A | N/A
20 | N/A | N/A | #VALUE! |  |  | 20 | N/A | N/A
21 | N/A | N/A | #VALUE! |  |  | 21 | N/A | N/A
22 | N/A | N/A | #VALUE! |  |  | 22 | N/A | N/A
23 | N/A | N/A | #VALUE! |  |  | 23 | N/A | N/A
24 | N/A | N/A | #VALUE! |  |  | 24 | N/A | N/A
Total | 200 | 2 |  |  |  | Total | 200 | 2
Comments (Attach separate sheet(s) as required.) | [9 empty] | Jesper Fredriksson
N/A | [8 empty] | Performed By:
2025 10 26
Date:
DO NOT PRINT BELOW
[7 empty] | Sample Size | Defects | Defect Rate
[7 empty] | Total | Total | Total Percent | CL
[6 empty] | 1 | 20 | 0 | 0% | 0.14863
[6 empty] | 2 | 40 | 0 | 0% | 0.090058
[6 empty] | 3 | 60 | 1 | 1.667% | 0.069411
[6 empty] | 4 | 80 | 1 | 1.25% | 0.058822
[6 empty] | 5 | 100 | 1 | 1% | 0.052345
[6 empty] | 6 | 120 | 2 | 1.667% | 0.049449
[6 empty] | 7 | 140 | 2 | 1.429% | 0.045947
[6 empty] | 8 | 160 | 2 | 1.25% | 0.043305
[6 empty] | 9 | 180 | 2 | 1.111% | 0.041237
[6 empty] | 10 | 200 | 2 | 1% | 0.03957
[6 empty] | 11 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
[6 empty] | 12 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
[6 empty] | 13 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
[6 empty] | 14 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
[6 empty] | 15 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
[6 empty] | 16 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
[6 empty] | 17 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
[6 empty] | 18 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
[6 empty] | 19 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
[6 empty] | 20 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
[6 empty] | 21 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
[6 empty] | 22 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
[6 empty] | 23 | #VALUE! | #VALUE! | #VALUE! | #VALUE!
[6 empty] | 24 | #VALUE! | #VALUE! | #VALUE! | #VALUE!

## Resample Data Summary
Data Summary
Part No. |  | GXMTBRIFULTRA50
Expected Lot Number |  | 92303
Expected Assay Name |  | Xpert MTB-RIF Ultra
First 10 Sample Type | 0 | No. Of Replicates | 10
Lot Check | VALID | Assay Name Check | VALID
No. Of Replicates Check | PASS | Expected Result | MTB NOT DETECTED
Cartridge ID | Result | Failure check | Pressure Check | Comment (drop-down menu available)
Sample Type | 0 | No. Of Replicates | 0
Lot Check | NO DATA | Assay Name Check | NO DATA
No. Of Replicates Check | FAIL | Expected Result | MTB NOT DETECTED
Cartridge ID | Result | Failure check | Pressure Check | Comment (drop-down menu available)
Sample Type | 1 | No. Of Replicates | 0
Lot Check | NO DATA | Assay Name Check | NO DATA
No. Of Replicates Check | FAIL | Expected Result | MTB DETECTED *;Rif Resistance DETECTED
Cartridge ID | Result | Failure check | Pressure Check | Comment (drop-down menu available)
Sample Type | 2 | No. Of Replicates | 0
Lot Check | VALID | Assay Name Check | NO DATA
No. Of Replicates Check | FAIL | Expected Result | MTB DETECTED *;Rif Resistance NOT DETECTED
Cartridge ID | Result | Failure check | Pressure Check | Comment (drop-down menu available)
Replaced | ¨N/A | No. Of Replicates | 0
Cartridge ID | Result | Failure check | Pressure Check | Comment (drop-down menu available)
Recorded By: |  | Date:

## Resample Mean Ct Evaluation
Sample Type 0 (Negative) |  |  |  |  | Sample Type 1 (TBMDR2) | [10 empty] | Sample Type 2 (TBWT)
Data Set | IS1081-IS6110 Ct | IS1081-IS6110 Ct Mean per Data Set | SPC Ct | SPC mean per Data Set | Data Set | IS1081-IS6110 Ct | rpo1 Ct | rpo2 Ct | rpo4 Ct | rpo1 Tm | rpo1 Tm Mean | rpo2 Tm | rpo2 Tm Mean | rpo4 Tm | rpo4 Tm Mean | Data Set | IS1081-IS6110 Ct | rpo1 Ct | rpo2 Ct | rpo3 Ct | rpo4 Ct | rpo1 Tm | rpo1 Tm Mean | rpo2 Tm | rpo2 Tm Mean | rpo3 Tm | rpo3 Tm Mean | rpo4 Tm | rpo4 Tm Mean
1 |  | #DIV/0! |  | #DIV/0! | 1 | [5 empty] | #DIV/0! |  | #DIV/0! |  | #DIV/0! | 1-5 | [6 empty] | #DIV/0! |  | #DIV/0! |  | #DIV/0! |  | #DIV/0!
[11 empty] | PASS |  | PASS |  | PASS
|  | #DIV/0! |  | #DIV/0! | [18 empty] | Not enough data |  | Not enough data |  | Not enough data |  | Not enough data
[5 empty] | 2 | [5 empty] | #DIV/0! |  | #DIV/0! |  | #DIV/0!
2 |  | #DIV/0! |  | #DIV/0! | 2 | [10 empty] | 6-10 | [6 empty] | #DIV/0! |  | #DIV/0! |  | #DIV/0! |  | #DIV/0!
2
[11 empty] | PASS |  | PASS |  | PASS
|  | #DIV/0! |  | #DIV/0! | [18 empty] | Not enough data |  | Not enough data |  | Not enough data |  | Not enough data
[5 empty] | 3 | [5 empty] | #DIV/0! |  | #DIV/0! |  | #DIV/0!
3 |  | #DIV/0! |  | #DIV/0! | 3 | [5 empty] | PASS |  | PASS |  | PASS |  | IS1081-IS6110 Ct Mean per Data Set | rpo1 Ct Mean per Data Set | rpo2 Ct Mean per Data Set | rpo3 Ct Mean per Data Set | rpo4 Ct Mean per Data Set
[17 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
3 | [16 empty] | Not enough data | Not enough data | Not enough data | Not enough data | Not enough data
[5 empty] | 4 | [5 empty] | #DIV/0! |  | #DIV/0! |  | #DIV/0!
|  | #DIV/0! |  | #DIV/0!
[11 empty] | PASS |  | PASS |  | PASS
4 |  | #DIV/0! |  | #DIV/0! | 4
[5 empty] | 5 | [5 empty] | #DIV/0! |  | #DIV/0! |  | #DIV/0!
4
|  | #DIV/0! |  | #DIV/0!
[11 empty] | PASS |  | PASS |  | PASS
5 |  | #DIV/0! |  | #DIV/0! | 6 | [5 empty] | #DIV/0! |  | #DIV/0! |  | #DIV/0!
[11 empty] | PASS |  | PASS |  | PASS
5 |  | #DIV/0! |  | #DIV/0!
[5 empty] | 7 | [5 empty] | #DIV/0! |  | #DIV/0! |  | #DIV/0!
6 |  | #DIV/0! |  | #DIV/0! | 6
6 | [10 empty] | PASS |  | PASS |  | PASS
|  | #DIV/0! |  | #DIV/0!
[5 empty] | 8 | [5 empty] | #DIV/0! |  | #DIV/0! |  | #DIV/0!
7 |  | #DIV/0! |  | #DIV/0! | 7 | [5 empty] | PASS |  | PASS |  | PASS
7 |  |  |  |  | 9 | [5 empty] | #DIV/0! |  | #DIV/0! |  | #DIV/0!
|  | #DIV/0! |  | #DIV/0!
[11 empty] | PASS |  | PASS |  | PASS
8 |  | #DIV/0! |  | #DIV/0! | 8
[5 empty] | 10 | [5 empty] | #DIV/0! |  | #DIV/0! |  | #DIV/0!
|  | #DIV/0! |  | #DIV/0!
8 | [10 empty] | PASS |  | PASS |  | PASS
9 |  | #DIV/0! |  | #DIV/0! |  | IS1081-IS6110 Ct Mean per Data Set | rpo1 Ct Mean per Data Set | rpo2 Ct Mean per Data Set | rpo4 Ct Mean per Data Set
[6 empty] | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
[6 empty] | Not enough data | Not enough data | Not enough data | Not enough data
9
|  | #DIV/0! |  | #DIV/0!
10 |  | #DIV/0! |  | #DIV/0!
10
|  | #DIV/0! |  | #DIV/0!

## Data - Resample Replaced
Assay | Assay Version | Sample ID | Cartridge S/N | Reagent Lot ID | Test Type | Instrument S/N | Module S/N | S/W Version | Start Time | Status | Test Result | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Target # | Analyte Result | Ct | EndPt | Curve Fit | Target Delta Ct | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | 2nd Deriv Peak Height | Cycles | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Analyte Name | Usage | Prb Chk 1 | Prb Chk 2 | Prb Chk 3 | Probe Check Result | Melt Peak Temperature | Melt Peak Height | Max Pressure (PSI) | Min Pressure (PSI) | Error | Patient ID | Quantitative Value | Quantitative Unit | Interpretation Flag | User

## Specifications
Sample Size |  |  | 210 | Resample Size | 200
Major Cartridge Integrity Test Failure Limits (PASS, RESAMPLE), Note: this specification is not in use. |  |  | 1 | 2
Maximum Pressure Limit |  |  | 90 | [6 empty] | Minor Functional - Indeterminate error code
MTB DETECTED*;Rif Resistance INDETERMINATE
Valid Sample Type Codes | Valid Replicate Check |  | Max No. of Single Minor Failure Mode for initial data | 4
0 | 100 |  | Max No. of Single Minor Failure Mode for initial+resample data | 10
1 | 80
First 10 | 10
2 | 20 |  | Sample Type 0 | Sample Type 1 | Sample Type 2
Part/Cat. No.* | Assay Name | Assay (P File) | Expected Result | Expected Result | Expected Result | Maximum No. Allowable Major and Minor Failures Total for PASS | Maximum No. Allowable Seal Test Failures Total for PASS | Maximum No. Allowable Major Functional Failures for PASS | Maximum No. Allowable Major Visual Inspection Failures for PASS | Maximum No. Allowable Major Failures Total for PASS | Maximum No. Allowable Major Functional Failures for RESAMPLE | Maximum No. Allowable Major Visual Inspection Failures for RESAMPLE | Maximum No. Allowable Major Failures Total for RESAMPLE | Sample Type for First 10 | Maximum No. Allowable Major and Minor Failures Total for PASS | Maximum No. Allowable Major Functional Failures for PASS | Maximum No. Allowable Major Visual Inspection Failures for PASS | Maximum No. Allowable Major Failures Total for PASS
GXMTBRIF-ULT-CN-10 | Xpert MTB-RIF Ultra | P0833 | MTB NOT DETECTED | MTB DETECTED *;Rif Resistance DETECTED | MTB DETECTED *;Rif Resistance NOT DETECTED | 9 | 9 | 1 | 1 | 1 | 3 | 3 | 3 | 0 | 10 | 5 | 5 | 5
GXMTBRIF-ULT-CN-50 | Xpert MTB-RIF Ultra | P0833 | MTB NOT DETECTED | MTB DETECTED *;Rif Resistance DETECTED | MTB DETECTED *;Rif Resistance NOT DETECTED | 9 | 9 | 1 | 1 | 1 | 3 | 3 | 3 | 0 | 10 | 5 | 5 | 5
GXMTBRIFULTRA10 | Xpert MTB-RIF Ultra | P0833 | MTB NOT DETECTED | MTB DETECTED *;Rif Resistance DETECTED | MTB DETECTED *;Rif Resistance NOT DETECTED | 9 | 9 | 1 | 1 | 1 | 3 | 3 | 3 | 0 | 10 | 5 | 5 | 5
GXMTBRIFULTRA50 | Xpert MTB-RIF Ultra | P0833 | MTB NOT DETECTED | MTB DETECTED *;Rif Resistance DETECTED | MTB DETECTED *;Rif Resistance NOT DETECTED | 9 | 9 | 1 | 1 | 1 | 3 | 3 | 3 | 0 | 10 | 5 | 5 | 5
GXMTB-ULTRA-HB-10 | Xpert MTB-RIF Ultra | P0833 | MTB NOT DETECTED | MTB DETECTED *;Rif Resistance DETECTED | MTB DETECTED *;Rif Resistance NOT DETECTED | 9 | 9 | 1 | 1 | 1 | 3 | 3 | 3 | 0 | 10 | 5 | 5 | 5
GXMTB-ULTRA-HB-50 | Xpert MTB-RIF Ultra | P0833 | MTB NOT DETECTED | MTB DETECTED *;Rif Resistance DETECTED | MTB DETECTED *;Rif Resistance NOT DETECTED | 9 | 9 | 1 | 1 | 1 | 3 | 3 | 3 | 0 | 10 | 5 | 5 | 5
GXMTBULTRAMII10 | Xpert MTB-RIF Ultra | P0833 | MTB NOT DETECTED | MTB DETECTED *;Rif Resistance DETECTED | MTB DETECTED *;Rif Resistance NOT DETECTED | 9 | 9 | 1 | 1 | 1 | 3 | 3 | 3 | 0 | 10 | 5 | 5 | 5
GXMTBULTRAMII50 | Xpert MTB-RIF Ultra | P0833 | MTB NOT DETECTED | MTB DETECTED *;Rif Resistance DETECTED | MTB DETECTED *;Rif Resistance NOT DETECTED | 9 | 9 | 1 | 1 | 1 | 3 | 3 | 3 | 0 | 10 | 5 | 5 | 5
PN7006537 | Xpert MTB-RIF Ultra | P0833 | MTB NOT DETECTED | MTB DETECTED *;Rif Resistance DETECTED | MTB DETECTED *;Rif Resistance NOT DETECTED | 9 | 9 | 1 | 1 | 1 | 3 | 3 | 3 | 0 | 10 | 5 | 5 | 5
PN7007774 | Xpert MTB-RIF Ultra | P0833 | MTB NOT DETECTED | MTB DETECTED *;Rif Resistance DETECTED | MTB DETECTED *;Rif Resistance NOT DETECTED | 9 | 9 | 1 | 1 | 1 | 3 | 3 | 3 | 0 | 10 | 5 | 5 | 5
PN7009436 | Xpert MTB-RIF Ultra | P0833 | MTB NOT DETECTED | MTB DETECTED *;Rif Resistance DETECTED | MTB DETECTED *;Rif Resistance NOT DETECTED | 9 | 9 | 1 | 1 | 1 | 3 | 3 | 3 | 0 | 10 | 5 | 5 | 5
PN9000779 | MTB-RIF Ultra v2 RUO | P0792 | MTB NOT DETECTED | MTB DETECTED *;Rif Resistance DETECTED | MTB DETECTED *;Rif Resistance NOT DETECTED | 9 | 9 | 1 | 1 | 1 | 3 | 3 | 3 | 0 | 10 | 5 | 5 | 5
* Note: List must be in ascending order per Excel ordering rules.
TYPE | Max SPC Ct | Max rpoB2 Ct | IS1081-IS6110 Ct | rpo1 Ct | rpo1 Tm | rpo2 Ct | rpo2 MUT Tm | rpo3 Ct | rpo3 Tm | rpo4 Ct | rpo4 MUT A Tm
0 | 28.2 | 28.4 | 39 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A
1 | N/A | N/A | 30.2 | 28.3 | 7 | 30 | 20 | N/A | N/A | 27.9 | 20
2 | N/A | N/A | 26.6 | 29.1 | 7 | 29.4 | 20 | 30.2 | 20 | 31.2 | 20

## HistoricalAccuStats
1 |  |  | 2 |  | 3 |  | 4 |  | 5 |  | 6 |  | 7 |  | 8 |  | 9 |  | 10 |  | 11 |  | 12 |  | 13 |  | 14 |  | 15 |  | 16 |  | 17 |  | 18 |  | 19 |  | 20 |  | 21 |  | 22 |  | 23 |  | 24
| 25 |  |  | 20 |  | 20 |  | 20 |  | 20 |  | 20 |  | 20 |  | 20 |  | 20 |  | 20 |  | #VALUE! |  | #VALUE! |  | #VALUE! |  | #VALUE! |  | #VALUE! |  | #VALUE! |  | #VALUE! |  | #VALUE! |  | #VALUE! |  | #VALUE! |  | #VALUE! |  | #VALUE! |  | #VALUE! |  | #VALUE!
Conf | 0.9 |  | Conf | 0.9 | Conf | 0.9 | Conf | 0.9 | Conf | 0.9 | Conf | 0.9 | Conf | 0.9 | Conf | 0.9 | Conf | 0.9 | Conf | 0.9 | Conf | 0.9 | Conf | 0.9 | Conf | 0.9 | Conf | 0.9 | Conf | 0.9 | Conf | 0.9 | Conf | 0.9 | Conf | 0.9 | Conf | 0.9 | Conf | 0.9 | Conf | 0.9 | Conf | 0.9 | Conf | 0.9 | Conf | 0.9
SS | 20 |  | SS | 40 | SS | 60 | SS | 80 | SS | 100 | SS | 120 | SS | 140 | SS | 160 | SS | 180 | SS | 200 | SS | #VALUE! | SS | #VALUE! | SS | #VALUE! | SS | #VALUE! | SS | #VALUE! | SS | #VALUE! | SS | #VALUE! | SS | #VALUE! | SS | #VALUE! | SS | #VALUE! | SS | #VALUE! | SS | #VALUE! | SS | #VALUE! | SS | #VALUE!
Def | 0.4 |  | Def | 0.8 | Def | 1.2 | Def | 1.6 | Def | 2 | Def | 2.4 | Def | 2.8 | Def | 3.2 | Def | 3.6 | Def | 4 | Def | #VALUE! | Def | #VALUE! | Def | #VALUE! | Def | #VALUE! | Def | #VALUE! | Def | #VALUE! | Def | #VALUE! | Def | #VALUE! | Def | #VALUE! | Def | #VALUE! | Def | #VALUE! | Def | #VALUE! | Def | #VALUE! | Def | #VALUE!
v1 | 2.8 |  | v1 | 3.6 | v1 | 4.4 | v1 | 5.2 | v1 | 6 | v1 | 6.8 | v1 | 7.6 | v1 | 8.4 | v1 | 9.2 | v1 | 10 | v1 | #VALUE! | v1 | #VALUE! | v1 | #VALUE! | v1 | #VALUE! | v1 | #VALUE! | v1 | #VALUE! | v1 | #VALUE! | v1 | #VALUE! | v1 | #VALUE! | v1 | #VALUE! | v1 | #VALUE! | v1 | #VALUE! | v1 | #VALUE! | v1 | #VALUE!
v2 | 39.2 |  | v2 | 78.4 | v2 | 117.6 | v2 | 156.8 | v2 | 196 | v2 | 235.2 | v2 | 274.4 | v2 | 313.6 | v2 | 352.8 | v2 | 392 | v2 | #VALUE! | v2 | #VALUE! | v2 | #VALUE! | v2 | #VALUE! | v2 | #VALUE! | v2 | #VALUE! | v2 | #VALUE! | v2 | #VALUE! | v2 | #VALUE! | v2 | #VALUE! | v2 | #VALUE! | v2 | #VALUE! | v2 | #VALUE! | v2 | #VALUE!
Est | 0.02 |  | Est | 0.02 | Est | 0.02 | Est | 0.02 | Est | 0.02 | Est | 0.02 | Est | 0.02 | Est | 0.02 | Est | 0.02 | Est | 0.02 | Est | #VALUE! | Est | #VALUE! | Est | #VALUE! | Est | #VALUE! | Est | #VALUE! | Est | #VALUE! | Est | #VALUE! | Est | #VALUE! | Est | #VALUE! | Est | #VALUE! | Est | #VALUE! | Est | #VALUE! | Est | #VALUE! | Est | #VALUE!
ARL | 50 |  | ARL | 50 | ARL | 50 | ARL | 50 | ARL | 50 | ARL | 50 | ARL | 50 | ARL | 50 | ARL | 50 | ARL | 50 | ARL | #VALUE! | ARL | #VALUE! | ARL | #VALUE! | ARL | #VALUE! | ARL | #VALUE! | ARL | #VALUE! | ARL | #VALUE! | ARL | #VALUE! | ARL | #VALUE! | ARL | #VALUE! | ARL | #VALUE! | ARL | #VALUE! | ARL | #VALUE! | ARL | #VALUE!
CI | 0.14863 |  | CI | 0.090058 | CI | 0.069411 | CI | 0.058822 | CI | 0.052345 | CI | 0.049449 | CI | 0.045947 | CI | 0.043305 | CI | 0.041237 | CI | 0.03957 | CI | #VALUE! | CI | #VALUE! | CI | #VALUE! | CI | #VALUE! | CI | #VALUE! | CI | #VALUE! | CI | #VALUE! | CI | #VALUE! | CI | #VALUE! | CI | #VALUE! | CI | #VALUE! | CI | #VALUE! | CI | #VALUE! | CI | #VALUE!
ARL | 6.7282 |  | ARL | 11.104 | ARL | 14.407 | ARL | 17.0005 | ARL | 19.1039 | ARL | 20.2227 | ARL | 21.7644 | ARL | 23.0919 | ARL | 24.2498 | ARL | 25.2717 | ARL | #VALUE! | ARL | #VALUE! | ARL | #VALUE! | ARL | #VALUE! | ARL | #VALUE! | ARL | #VALUE! | ARL | #VALUE! | ARL | #VALUE! | ARL | #VALUE! | ARL | #VALUE! | ARL | #VALUE! | ARL | #VALUE! | ARL | #VALUE! | ARL | #VALUE!
1 |  |  | 2 |  | 3 |  | 4 |  | 5 |  | 6 |  | 7 |  | 8 |  | 9 |  | 10 |  | 11 |  | 12 |  | 13 |  | 14 |  | 15 |  | 16 |  | 17 |  | 18 |  | 19 |  | 20 |  | 21 |  | 22 |  | 23 |  | 24
F-statistic |  |  | F-statistic |  | F-statistic |  | F-statistic |  | F-statistic |  | F-statistic |  | F-statistic |  | F-statistic |  | F-statistic |  | F-statistic |  | F-statistic |  | F-statistic |  | F-statistic |  | F-statistic |  | F-statistic |  | F-statistic |  | F-statistic |  | F-statistic |  | F-statistic |  | F-statistic |  | F-statistic |  | F-statistic |  | F-statistic |  | F-statistic
2.444 |  |  | 2.1554 |  | 1.9935 |  | 1.8846 |  | 1.8044 |  | 1.7993 |  | 1.7388 |  | 1.6899 |  | 1.6494 |  | 1.615 |  | #VALUE! |  | #VALUE! |  | #VALUE! |  | #VALUE! |  | #VALUE! |  | #VALUE! |  | #VALUE! |  | #VALUE! |  | #VALUE! |  | #VALUE! |  | #VALUE! |  | #VALUE! |  | #VALUE! |  | #VALUE!
| cl | est
1 | 0.14863 | 0.0018
2 | 0.090058 | 0.0018
3 | 0.069411 | 0.0018
4 | 0.058822 | 0.0018
5 | 0.052345 | 0.0018
6 | 0.049449 | 0.0018
7 | 0.045947 | 0.0018
8 | 0.043305 | 0.0018
9 | 0.041237 | 0.0018
10 | 0.03957 | 0.0018
11 | #VALUE! | 0.0018
12 | #VALUE! | 0.0018
13 | #VALUE! | 0.0018
14 | #VALUE! | 0.0018
15 | #VALUE! | 0.0018
16 | #VALUE! | 0.0018
17 | #VALUE! | 0.0018
18 | #VALUE! | 0.0018
19 | #VALUE! | 0.0018
20 | #VALUE! | 0.0018
21 | #VALUE! | 0.0018
22 | #VALUE! | 0.0018
23 | #VALUE! | 0.0018
24 | #VALUE! | 0.0018

## For Reference Only
Reasons for replacements | Comments for Valid Test Replicates | NC No. | [7 empty] | Type Of Resample
[7 empty] | NA | NA | YES
Instrument error replaced per D12547 Rev __ | See Replaced Data | __NC____ | See Replaced Data |  |  | __NC____ | __LFI__ | YES | NO | Visual Only Resample
Replaced per __LFI___ | NA | NA | NA |  |  | NA |  | NO | Resample Needed | Functional and Visual Resample
NA | Bent reaction tube | [8 empty] | Delam and Visual Resample
| Broken/missing reaction tube | [8 empty] | NA
Cracked chimneys
| Desiccant in center plunger hole |  | x(Seal Test Only) | x | NA
| Film delamination |  | x | NA | PASS
| Fluids/crystals on the top/inside of the cartridge |  | NA |  | RESAMPLE
Incorrect lid
Label incorrect/blurry
Leaky cartridge
Missing or noticeably under or overfilled on board reagent